## report
ID | Department | Name | Instrument ID | Instrument Breakdown Date | Contact Time | Suport Time | Complete Time | Type | Detail | Action | Offline/ Online Support | Suport Persion | Contact Persion | Current Status | Note | Test affected during breakdown (yes/No) | Name of the test affected | No of test affected | Image | File
46288622 | GLP | Tube Assecssment Module | TAM | 2023-08-09 11:38 | 2023-08-09 11:38 | 2023-08-09 11:38 | 2023-08-09 11:45 | Breakdown | Finger's rubber torned | Change finger | Offline | BME | BME | Done |  | FALSE |  |  | report_Images/46288622.Image.044707.jpg
a733af57 | GLP | Cars and Track | CTR | 2023-08-04 15:45 | 2023-08-04 15:45 | 2023-08-04 15:45 | 2023-08-04 16:25 | Maintenance | Prevent Maintenance | Maintenance | Offline | BME | BME | Done |  | FALSE
f46db097 | Molecular Biology | KingFisher Presto | KFP-01 | 2023-08-02 15:09 | 2023-08-02 15:09 | 2023-08-02 15:12 | 2023-08-02 15:49 | Breakdown | Heater out of possition | Handle to keep going process; clean axis, repower device | Offline | BME | BME | Done |  | FALSE
7fb5063d | GLP | Cars and Track | CTR | 2023-08-01 14:10 | 2023-08-01 14:10 | 2023-08-01 14:10 | 2023-08-01 15:00 | Breakdown | Nodes disconnected | Check segment, connecting cables; repower, reconnect | Online | Mr. Hoàng | BME | Follow up | Need check log | FALSE |  |  | report_Images/7fb5063d.Image.080429.jpg
63ca6d1b | GLP | Tube Assecssment Module | TAM | 2023-07-28 11:57 | 2023-07-28 11:57 | 2023-07-28 11:59 | 2023-07-28 12:11 | Breakdown | Finger cut through tube cap | Take out, clean, re initialize | Offline | BME | BME | Done |  | FALSE |  |  | report_Images/63ca6d1b.Image.051649.jpg
31 | GLP | Cars and Track | CTR | 2023-07-28 11:00 | 2023-07-28 11:00 | 2023-07-28 11:00 | 2023-07-28 12:00 | Maintenance | Prevent maintenance | Maintenance | Offline | BME |  | Done |  | FALSE
9d704970 | GLP | SAL User Interface | LAS-UI1 | 2023-07-27 09:00 | 2023-07-27 09:00 | 2023-07-27 09:00 | 2023-07-27 10:00 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Hoàng | BME | Done |  | FALSE |  |  | report_Images/9d704970.Image.093305.png | report_Files_/9d704970.File.093305.pdf
cf93008a | GLP | SAL User Interface | LAS-UI2 | 2023-07-27 09:00 | 2023-07-27 09:00 | 2023-07-27 09:00 | 2023-07-27 10:00 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Hoàng | BME | Done |  | FALSE |  |  | report_Images/cf93008a.Image.093735.png | report_Files_/cf93008a.File.093735.pdf
d0ebf150 | GLP | SAL User Interface | LAS-UI3 | 2023-07-27 09:00 | 2023-07-27 09:00 | 2023-07-27 09:00 | 2023-07-27 10:00 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Hoàng | BME | Done |  | FALSE |  |  | report_Images/d0ebf150.Image.093414.png | report_Files_/d0ebf150.File.093414.pdf
e7492a43 | GLP | Rack Builder Module | RBM-C8000 | 2023-07-27 09:00 | 2023-07-27 09:00 | 2023-07-27 09:00 | 2023-07-27 10:00 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Hoàng | BME | Done |  | FALSE |  |  | report_Images/e7492a43.Image.093613.png | report_Files_/e7492a43.File.093613.pdf
076608ae | GLP | Storage on SM | STORE-01 | 2023-07-27 09:00 | 2023-07-27 09:00 | 2023-07-27 09:00 | 2023-07-27 10:00 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Hoàng | BME | Done |  | FALSE |  |  | report_Images/076608ae.Image.093101.png | report_Files_/076608ae.File.093101.pdf
0be204a6 | GLP | Tosho Connect Module | TOS-01 | 2023-07-27 09:00 | 2023-07-27 09:00 | 2023-07-27 09:00 | 2023-07-27 10:00 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Hoàng | BME | Done |  | FALSE |  |  | report_Images/0be204a6.Image.090323.png | report_Files_/0be204a6.File.090323.pdf
756ddbe6 | GLP | Bulk Loader Module | BLM | 2023-07-26 09:00 | 2023-07-26 09:00 | 2023-07-26 09:00 | 2023-07-26 10:00 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Hoàng | BME | Done |  | FALSE |  |  | report_Images/756ddbe6.Image.092224.png | report_Files_/756ddbe6.File.092224.pdf
b601195e | GLP | Decaper Module Double | DM-01 | 2023-07-26 09:00 | 2023-07-26 09:00 | 2023-07-26 09:00 | 2023-07-26 10:00 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Hoàng | BME | Done |  | FALSE |  |  | report_Images/b601195e.Image.092601.png | report_Files_/b601195e.File.092601.pdf
f7087bf7 | GLP | Tempus Module | IOM_TIO-01 | 2023-07-26 09:00 | 2023-07-26 09:00 | 2023-07-26 09:00 | 2023-07-26 10:00 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Hoàng | BME | Done |  | FALSE |  |  | report_Images/f7087bf7.Image.091733.png | report_Files_/f7087bf7.File.091733.pdf
133cba5e | GLP | Tempus Module | IOM_TIO-02 | 2023-07-26 09:00 | 2023-07-26 09:00 | 2023-07-26 09:00 | 2023-07-26 10:00 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Hoàng | BME | Done |  | FALSE |  |  | report_Images/133cba5e.Image.092103.png | report_Files_/133cba5e.File.092103.pdf
78d7cb02 | GLP | Loader Module | LM-01 | 2023-07-26 09:00 | 2023-07-26 09:00 | 2023-07-26 09:00 | 2023-07-26 10:00 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Hoàng | BME | Done |  | FALSE |  |  | report_Images/78d7cb02.Image.091922.png | report_Files_/78d7cb02.File.091922.pdf
9278bdc6 | GLP | Remover Cap Module | REM-01 | 2023-07-26 09:00 | 2023-07-26 09:00 | 2023-07-26 09:00 | 2023-07-26 10:00 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Hoàng | BME | Done |  | FALSE |  |  | report_Images/9278bdc6.Image.092943.png | report_Files_/9278bdc6.File.092943.pdf
b73351a2 | GLP | Recapper Module | RM-01 | 2023-07-26 09:00 | 2023-07-26 09:00 | 2023-07-26 09:00 | 2023-07-26 10:00 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Hoàng | BME | Done |  | FALSE |  |  | report_Images/b73351a2.Image.092813.png | report_Files_/b73351a2.File.092813.pdf
866fe204 | AO | Tempus | Tempus600 | 2023-07-25 10:51 | 2023-07-25 10:52 | 2023-07-25 10:53 | 2023-07-25 11:10 | Breakdown | Stuck sample in break | Take out and re-intialization | Offline | BME | BME | Follow up |  | FALSE
6e098c95 | GLP | Aliquoter Module | AM-01 | 2023-07-25 09:00 | 2023-07-25 09:00 | 2023-07-25 09:00 | 2023-07-25 10:00 | Maintenance | Prevent Maintenance |  | Offline | Mr. Hoàng | BME | Done |  | FALSE |  |  | report_Images/6e098c95.Image.102002.png | report_Files_/6e098c95.File.102002.pdf
8e91e0f2 | GLP | Centrifuge Module | CM-01 | 2023-07-25 09:00 | 2023-07-25 09:00 | 2023-07-25 09:00 | 2023-07-25 10:00 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Hoàng | BME | Done |  | FALSE |  |  | report_Images/8e91e0f2.Image.092324.png | report_Files_/8e91e0f2.File.092324.pdf
88539b08 | GLP | Centrifuge Module | CM-02 | 2023-07-25 09:00 | 2023-07-25 09:00 | 2023-07-25 09:00 | 2023-07-25 10:00 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Hoàng | BME | Done |  | FALSE |  |  | report_Images/88539b08.Image.091014.png
64b5d1a8 | GLP | Centrifuge Module | CM-03 | 2023-07-25 09:00 | 2023-07-25 09:00 | 2023-07-25 09:00 | 2023-07-25 10:00 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Hoàng | BME | Done |  | FALSE |  |  | report_Images/64b5d1a8.Image.090916.png | report_Files_/64b5d1a8.File.090916.pdf
097a077c | GLP | Tube Assecssment Module | TAM | 2023-07-25 09:00 | 2023-07-25 09:00 | 2023-07-25 09:00 | 2023-07-25 10:00 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Hoàng | BME | Done |  | FALSE |  |  | report_Images/097a077c.Image.091350.png | report_Files_/097a077c.File.091350.pdf
24ba6eb2 | GLP | Stramax Connect Module | STR-01 | 2023-07-24 13:52 | 2023-07-24 14:02 | 2023-07-24 14:32 | 2023-07-24 15:32 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Hoàng | BME | Done |  | FALSE |  |  |  | report_Files_/24ba6eb2.File.083320.pdf
b1f05e86 | GLP | Input/ Output Module | IOM-03 | 2023-07-24 13:41 | 2023-07-24 13:51 | 2023-07-24 14:21 | 2023-07-24 15:21 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Hoàng | BME | Done |  | FALSE |  |  | report_Images/b1f05e86.Image.082340.png | report_Files_/b1f05e86.File.083104.pdf
700858bc | GLP | Buffer Module | BM-01 | 2023-07-24 08:30 | 2023-07-24 08:30 | 2023-07-24 08:30 | 2023-07-24 10:30 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Hoàng | BME | Done |  | FALSE |  |  | report_Images/700858bc.Image.081807.png
3d4e650c | GLP | Buffer Module | BM-02 | 2023-07-24 08:30 | 2023-07-24 08:30 | 2023-07-24 08:30 | 2023-07-24 10:00 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Hoàng | BME | Done |  | FALSE |  |  | report_Images/3d4e650c.Image.081555.png | report_Files_/3d4e650c.File.083736.pdf
05dd9c15 | ELISA | DAR800 | ELISA-02 | 2023-07-21 14:10 | 2023-07-21 14:10 | 2023-07-21 14:20 | 2023-07-21 16:00 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Bình | BME | Done |  | FALSE
3fcc67b3 | Microbiology | HeliFan Plus C13 | HP-01 | 2023-07-21 14:10 | 2023-07-21 14:10 | 2023-07-21 14:20 | 2023-07-21 16:00 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Vinh | BME | Done |  | FALSE
ccb66594 | ELISA | DAW50 | ELISA-01 | 2023-07-21 14:10 | 2023-07-21 14:10 | 2023-07-21 14:20 | 2023-07-21 15:52 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Bình | BME | Done |  | FALSE
7bb4ace6 | ELISA | DAW50TS | ELISA-03 | 2023-07-21 14:10 | 2023-07-21 14:10 | 2023-07-21 14:20 | 2023-07-21 15:52 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Bình | BME | Done |  | FALSE
78583fa6 | GLP | Input/ Output Module | IOM-02 | 2023-07-21 13:57 | 2023-07-21 14:07 | 2023-07-21 14:37 | 2023-07-21 15:37 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Hoàng | BME | Done |  | FALSE |  |  |  | report_Files_/78583fa6.File.083801.pdf
a4ec4c6d | GLP | Input/ Output Module | IOM-02 | 2023-07-21 13:55 | 2023-07-21 14:05 | 2023-07-21 14:35 | 2023-07-21 15:35 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Hoàng | BME | Done |  | FALSE |  |  |  | report_Files_/a4ec4c6d.File.083539.pdf
8e88b205 | GLP | Tube Assecssment Center | TAC2 | 2023-07-21 13:53 | 2023-07-21 14:03 | 2023-07-21 14:33 | 2023-07-21 15:33 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Hoàng | BME | Done |  | FALSE |  |  |  | report_Files_/8e88b205.File.083448.pdf
957b4455 | Hematology - Coagulation | STA R MAX 3 | DM-03 | 2023-07-20 15:50 | 2023-07-20 16:00 | 2023-07-20 16:05 | 2023-07-20 16:30 | Breakdown | Stucked cuvette | Clean cuvette pusher | Online | Mr. Huy | BME | Done |  | FALSE
7a1da242 | Cyto Lab | FUME HOOD | FH-01 | 2023-07-18 16:56 | 2023-07-18 16:56 | 2023-07-18 16:56 | 2023-07-18 17:36 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Tài | BME | Done |  | FALSE
eac7e26e | Cyto Lab | OVEN | OVEN-01 | 2023-07-18 16:56 | 2023-07-18 16:56 | 2023-07-18 16:56 | 2023-07-18 17:36 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Tài | BME | Done |  | FALSE
c19e1b3c | Bio-Immuno | Alinity C | AC-2 | 2023-07-18 14:01 | 2023-07-18 14:02 | 2023-07-18 14:05 | 2023-07-18 14:20 | Breakdown | Optic error | Change lamp | Online | Mr. Vũ | BME | Done |  | FALSE
34513 | GLP | Cars and Track | CTR | 2023-07-14 14:00 | 2023-07-14 14:00 | 2023-07-14 14:00 | 2023-07-14 15:00 | Maintenance | Prevent maintenance | Maintenance | Offline | BME |  | Done |  | FALSE
88b44636 | AO | Tempus | Tempus600 | 2023-07-14 08:15 | 2023-07-14 08:18 | 2023-07-14 08:48 | 2023-07-14 09:48 | Breakdown | Tube not send on Carousel | Clean sensor, carousel | Offline | BME | BME | Done |  | FALSE
78dd6b7a | All | Sample Ladder | Lift | 2023-07-13 13:46 | 2023-07-13 13:56 | 2023-07-13 14:26 | 2023-07-13 15:26 | Breakdown | counterweight rope break | Connect cable | Offline | Mr. Lâm | BME | Done |  | FALSE
a30d16c6 | Bio-Immuno | Alinity I | AI-2 | 2023-07-13 11:43 | 2023-07-13 11:53 | 2023-07-13 12:00 | 2023-07-13 12:23 | Breakdown | Board communication error | Change fuse | Offline | Mr. Hoàng | BME | Done |  | FALSE
86aa6b27 | AO | Tempus | Tempus600 | 2023-07-12 16:27 | 2023-07-13 10:37 | 2023-07-13 11:07 | 2023-07-13 12:07 | Breakdown | Sending sensor error | Clean sensor and run test Programme | Offline | Mr. Hoàng | BME | Follow up |  | FALSE |  |  | report_Images/86aa6b27.Image.050912.jpg
c3a25f1f | Bio-Immuno | Alinity I | AI-1 | 2023-07-08 09:43 | 2023-07-08 09:45 | 2023-07-08 10:15 | 2023-07-08 10:43 | Breakdown | Pre-trigger level sensor error | Change level sensor | Offline | Mr. Tuấn Anh | Ms. Trân | Follow up |  | FALSE
46 | GLP | Cars and Track | CTR | 2023-07-07 14:00 | 2023-07-07 14:00 | 2023-07-07 14:00 | 2023-07-07 15:00 | Maintenance | Prevent maintenance | Maintenance | Offline | BME |  | Done |  | FALSE
13f9b0bf | GLP | Centrifuge Module | CM-01 | 2023-07-06 15:00 | 2023-07-06 15:01 | 2023-07-06 15:05 | 2023-07-06 15:15 | Breakdown | Finger cut through tube cap | Take out all tubes; clean and re init machine | Offline | BME |  | Done |  | FALSE |  |  | report_Images/13f9b0bf.Image.081909.jpg
2fb942ae | Hematology - Coagulation | Tosoh HLC-723 G11 LA analyzer | HbA1c-02 | 2023-07-06 10:07 | 2023-07-06 10:17 | 2023-07-06 10:47 | 2023-07-06 11:47 | Breakdown | Valve have error | Check liquid line, valve and restart | Online | Mr. Trường | Mr. Lâm | Follow up |  | FALSE
ef83697a | Hematology - Coagulation | Alinity H-series | Hq-1 | 2023-07-03 11:29 | 2023-07-03 11:39 | 2023-07-03 12:09 | 2023-07-03 13:09 | Breakdown | WBC not normal | Clean WBC tank | Offline | Mr. Tới | BME | Follow up |  | FALSE
65 | GLP | Cars and Track | CTR | 2023-06-30 17:00 | 2023-06-30 17:00 | 2023-06-30 17:00 | 2023-06-30 18:00 | Maintenance | Prevent maintenance | Maintenance | Offline | BME |  | Done |  | FALSE
56456 | GLP | Cars and Track | CTR | 2023-06-23 17:00 | 2023-06-23 17:00 | 2023-06-23 17:00 | 2023-06-23 18:00 | Maintenance | Prevent maintenance | Maintenance | Offline | BME |  | Done |  | FALSE
22e5e5af | Hematology - Coagulation | Alinity H-series | Hq-1 | 2023-06-23 14:52 | 2023-06-23 15:02 | 2023-06-23 15:32 | 2023-06-23 16:32 | Breakdown | WBC low | Clean and check again | Offline | Mr. Trúc | Mr. Lâm | Done |  | TRUE | QC | 6
e7cc2366 | Hematology - Coagulation | Alinity H-series | Hq-2 | 2023-06-23 14:01 | 2023-06-23 14:11 | 2023-06-23 14:41 | 2023-06-23 15:41 | Breakdown | RBC QC fail | Change RBC factor | Offline | Ms. Phương | BME | Follow up |  | FALSE
46efea6f | Molecular Biology | Agilent Microplate Centrifuge | MC-01 | 2023-06-23 13:06 | 2023-06-23 13:16 | 2023-06-23 13:46 | 2023-06-23 14:46 | Maintenance | Prevent maintenance | Maintenance | Offline | BME | Ms. Nhân | Done |  | FALSE
d6051536 | Molecular Biology | Microplate Washer | MPW-01 | 2023-06-23 13:06 | 2023-06-23 13:16 | 2023-06-23 13:46 | 2023-06-23 14:46 | Maintenance | Prevent maintenance | Maintenance | Offline | BME | Ms. Nhân | Done |  | FALSE
b10b4ace | Molecular Biology | Automated Laboratory Plate Sealer | ALPS | 2023-06-23 13:05 | 2023-06-23 13:15 | 2023-06-23 13:45 | 2023-06-23 14:45 | Maintenance | Prevent maintenance | Maintenance | Offline | BME | Ms. Nhân | Done |  | FALSE
0fb70e92 | Molecular Biology | Bentrack | Ben | 2023-06-23 13:05 | 2023-06-23 13:15 | 2023-06-23 13:45 | 2023-06-23 14:45 | Maintenance | Prevent maintenance | Maintenance | Offline | BME | Ms. Nhân | Done |  | FALSE
034afa29 | Molecular Biology | Spinnaker Microplate Robot | Spinnaker | 2023-06-23 13:05 | 2023-06-23 13:15 | 2023-06-23 13:45 | 2023-06-23 14:45 | Maintenance | Prevent maintenance | Maintenance | Offline | BME | Ms. Nhân | Done |  | FALSE
013ef831 | Molecular Biology | KingFisher Presto | KFP-03 | 2023-06-23 13:04 | 2023-06-23 13:14 | 2023-06-23 13:44 | 2023-06-23 14:44 | Maintenance | Prevent maintenance | Maintenance | Offline | BME | Ms. Nhân | Done |  | FALSE
0ae59af1 | Molecular Biology | KingFisher Presto | KFP-04 | 2023-06-23 13:04 | 2023-06-23 13:14 | 2023-06-23 13:44 | 2023-06-23 14:44 | Maintenance | Prevent maintenance | Maintenance | Offline | BME | Ms. Nhân | Done |  | FALSE
8ae508fe | Molecular Biology | MultiDrop ComBI | MDC-01 | 2023-06-23 13:04 | 2023-06-23 13:14 | 2023-06-23 13:44 | 2023-06-23 14:44 | Maintenance | Prevent maintenance | Maintenance | Offline | BME | Ms. Nhân | Done |  | FALSE
ab0922e9 | Molecular Biology | MultiDrop ComBI | MDC-02 | 2023-06-23 13:04 | 2023-06-23 13:14 | 2023-06-23 13:44 | 2023-06-23 14:44 | Maintenance | Prevent maintenance | Maintenance | Offline | BME | Ms. Nhân | Done |  | FALSE
c0c02053 | Molecular Biology | MultiDrop ComBI | MDC-03 | 2023-06-23 13:04 | 2023-06-23 13:14 | 2023-06-23 13:44 | 2023-06-23 14:44 | Maintenance | Prevent maintenance | Maintenance | Offline | BME | Ms. Nhân | Done |  | FALSE
92bf04f8 | Molecular Biology | KingFisher Presto | KFP-02 | 2023-06-23 13:03 | 2023-06-23 13:13 | 2023-06-23 13:43 | 2023-06-23 14:43 | Maintenance | Prevent maintenance | Maintenance | Offline | BME | Ms. Nhân | Done |  | FALSE
3957ffa7 | Molecular Biology | KingFisher Presto | KFP-01 | 2023-06-23 12:59 | 2023-06-23 13:09 | 2023-06-23 13:39 | 2023-06-23 14:39 | Maintenance | Prevent maintenance | Maintenance | Offline | BME | Ms. Nhân | Done |  | FALSE
8b54985c | Bio-Immuno | Cobas e801 Analytical unit | MD-09 | 2023-06-23 07:50 | 2023-06-23 08:00 | 2023-06-23 09:30 | 2023-06-23 10:30 | Breakdown | Reagent stuck | Take out and re-init reagent path | Offline | Mr. Khang | BME | Follow up |  | FALSE
67924cc0 | GLP | Tempus Module | IOM_TIO-01 | 2023-06-22 14:41 | 2023-06-22 14:51 | 2023-06-22 15:21 | 2023-06-22 16:21 | Breakdown | Change break part | Change new version of break | Offline | Mr. Hoàng | BME | Follow up | Test | FALSE |  |  | report_Images/67924cc0.Image.092258.jpg
26e1ae55 | Hematology - Coagulation | Alinity H-series | Hq-2 | 2023-06-21 14:03 | 2023-06-21 14:13 | 2023-06-21 14:43 | 2023-06-21 15:43 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Tới | BME | Done |  | FALSE
093337c5 | Hematology - Coagulation | Alinity H-series | Hq-1 | 2023-06-21 14:02 | 2023-06-21 14:12 | 2023-06-21 14:42 | 2023-06-21 15:42 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Tới | BME | Done |  | FALSE
4dd9e8ed | Bio-Immuno | RO system Roche | RO-MB | 2023-06-21 12:16 | 2023-06-21 12:26 | 2023-06-21 13:56 | 2023-06-21 14:56 | Maintenance | Prevent maintenance | Change filter | Offline | Mr. Tín | BME | Done |  | FALSE
f7b59e30 | All | Sample Ladder | Lift | 2023-06-20 09:33 | 2023-06-20 09:43 | 2023-06-20 10:13 | 2023-06-20 11:13 | Breakdown | Can't move | Check and reset sensor | Offline | BG | BME | Done |  | FALSE
dbb2ccf0 | Bio-Immuno | Alinity I | AI-2 | 2023-06-20 07:00 | 2023-06-20 07:10 | 2023-06-20 07:20 | 2023-06-20 08:40 | Breakdown | Sample probe bend | Change probe | Offline | Mr. Tới | BME | Done |  | FALSE
3703345a | Hematology - Coagulation | Alinity H-series | Hq-2 | 2023-06-19 14:46 | 2023-06-19 14:56 | 2023-06-19 15:26 | 2023-06-19 16:26 | Breakdown | Fluid entered buffer tank | Change valve, clean buffer tank | Offline | Mr. Tới | Mr. Trung | Done |  | TRUE | QC | 10 | report_Images/3703345a.Image.092938.jpg
354 | GLP | Cars and Track | CTR | 2023-06-16 09:44 | 2023-06-16 09:44 | 2023-06-16 09:45 | 2023-06-16 11:14 | Maintenance | Prevent maintenance | Maintenance | Offline | BME |  | Done |  | FALSE
41175c07 | Hematology - Coagulation | Tosoh HLC-723 G11 analyzer | HbA1c-01 | 2023-06-15 15:45 | 2023-06-15 15:55 | 2023-06-15 16:25 | 2023-06-15 17:25 | Maintenance | Prevent maintenance | Maintenance | Offline | Mr. Trường | BME | Follow up |  | FALSE |  |  | report_Images/41175c07.Image.102626.jpg
af9dde84 | Hematology - Coagulation | Tosoh HLC-723 G11 LA analyzer | HbA1c-02 | 2023-06-15 15:43 | 2023-06-15 15:53 | 2023-06-15 16:23 | 2023-06-15 17:23 | Maintenance | Prevent maintenance | Maintenance | Offline | Mr. Trường | BME | Follow up |  | FALSE |  |  | report_Images/af9dde84.Image.102513.jpg
6a08fc90 | Hematology - Coagulation | Alinity H-series | Hq-2 | 2023-06-15 09:35 | 2023-06-15 09:45 | 2023-06-15 10:00 | 2023-06-15 16:15 | Breakdown | QC WBC out of range | Clean dilution cup, check tubing, run auto clean, run QC | Offline | Mr. Tới | BME | Follow up | Pass QC | FALSE |  |  | report_Images/6a08fc90.Image.091843.jpg
4ab0d176 | Microbiology | Previ color gram | GR-01 | 2023-06-15 13:46 | 2023-06-15 13:56 | 2023-06-15 14:00 | 2023-06-15 14:30 | Maintenance | Prevent maintenance | Maintenance | Offline | Mr. Trường | BME | Follow up |  | FALSE |  |  | report_Images/4ab0d176.Image.102820.jpg
d736ce8e | GLP | Tempus Module | IOM_TIO-01 | 2023-06-12 09:36 | 2023-06-12 09:46 | 2023-06-12 10:16 | 2023-06-12 11:16 | Breakdown | Change rubber | Change Break's Rubber | Offline | BME |  | Done |  | FALSE
e0f2f4d9 | GLP | Tempus Module | IOM_TIO-02 | 2023-06-12 09:35 | 2023-06-12 09:45 | 2023-06-12 10:15 | 2023-06-12 11:15 | Breakdown | Change rubber | Change Break's Rubber | Offline | BME |  | Done |  | FALSE
3ce348cc | GLP | Cars and Track | CTR | 2023-06-12 09:44 | 2023-06-12 09:44 | 2023-06-12 09:45 | 2023-06-12 11:14 | Maintenance | Prevent maintenance | Maintenance | Offline | BME |  | Done |  | FALSE
57aa0b6d | Bio-Immuno | Cobas e801 Analytical unit | MD-09 | 2023-06-07 12:25 | 2023-06-07 12:30 | 2023-06-07 13:00 | 2023-06-07 14:15 | Breakdown | Barcode read failed | Check and clean barcode reader; check sample barcode --> have error on barcode | Offline | Mr. Khang | BME | Follow up | Need to change barcode roll | FALSE |  |  | report_Images/57aa0b6d.Image.041657.png
c056e822 | GLP | Decaper Module Double | DM-01 | 2023-06-05 08:10 | 2023-06-05 08:10 | 2023-06-05 08:30 | 2023-06-06 14:50 | Breakdown | Car stuck on GLP --> AP failed, cut Car pins | Check AP, change car pins | Offline | Mr. Trúc, Mr. Hoàng | BME | Done |  | FALSE
3ce348cc | GLP | Cars and Track | CTR | 2023-06-02 14:44 | 2023-06-02 14:44 | 2023-06-02 14:45 | 2023-06-02 15:14 | Maintenance | Prevent maintenance | Maintenance | Offline | BME |  | Done
55a39892 | Hematology - Coagulation | Alinity H-series | Hq-1 | 2023-06-02 08:35 | 2023-06-02 08:38 | 2023-06-02 09:05 | 2023-06-02 10:15 | Breakdown | Result failed | Change sample needle | Online | Mr. Tới | BME | Done |  | FALSE
1978e36a | Bio-Immuno | RO system Abbott | RO-HP | 2023-05-30 14:08 | 2023-05-30 14:10 | 2023-05-30 14:30 | 2023-05-30 16:48 | Breakdown | Pressure pump have error, Prevent Maintenance | Change pump, filters | Offline | Mr. Tâm | BME | Done |  | FALSE |  |  | report_Images/1978e36a.Image.035159.png
dbaa31ba | Hematology - Coagulation | Automated Blood typing machine | ABO-01 | 2023-05-29 14:30 | 2023-05-29 02:30 | 2023-05-29 14:32 | 2023-05-29 16:51 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Điệu | BME | Done |  | FALSE
407623c2 | Hematology - Coagulation | Automated Blood typing machine | ABO-01 | 2023-05-29 14:04 | 2023-05-29 14:10 | 2023-05-29 14:45 | 2023-05-29 16:15 | Maintenance | Prevent Maintenance | Clean; Check and calibration camera; Run Sample | Offline | Mr. Điệu | BME | Done |  | FALSE |  |  | report_Images/407623c2.Image.034716.png
f9302257 | GLP | Tempus Module | IOM_TIO-01 | 2023-05-27 09:00 | 2023-05-27 09:01 | 2023-05-27 10:11 | 2023-05-27 10:45 | Breakdown | Gripper finger broke | Change finger | Online | Mr. Hòang | BME | Done |  | FALSE |  |  | report_Images/f9302257.Image.035032.jpg
27053779 | Bio-Immuno | Cobas e801 Analytical unit | MD-09 | 2023-05-26 13:00 | 2023-05-26 13:05 | 2023-05-26 14:30 | 2023-05-26 16:45 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Khang | BME | Done |  | FALSE |  |  | report_Images/27053779.Image.044700.png
0f970eff | GLP | Recapper Module | RM-01 | 2023-05-23 14:00 | 2023-05-23 14:03 | 2023-05-24 14:06 | 2023-05-26 14:20 | Breakdown | Robot crash detected (Robot 1) | Cleaned and lubricated recap head -> still failed -> change RM griper | Offline | Mr. Hoàng | BME | Done | 803012411 24/05/2023 5210007482 24/05/2023 10 A05SFG.41365.01 26344503 RM GRIPPER UNIT S A44 1 | FALSE |  |  | report_Images/0f970eff.Image.035855.png
fc74f7ad | Bio-Immuno | Cobas8000 CORE MODULE | Core 4 | 2023-05-26 13:00 | 2023-05-26 13:00 | 2023-05-26 13:15 | 2023-05-26 14:15 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Khang | BME | Done |  | FALSE |  |  | report_Images/fc74f7ad.Image.044323.png
3ce348cc | GLP | Cars and Track | CTR | 2023-05-26 09:44 | 2023-05-26 09:44 | 2023-05-26 09:45 | 2023-05-26 11:14 | Maintenance | Prevent maintenance | Maintenance | Offline | BME |  | Done
80e0c740 | Hematology - Coagulation | Alinity H-series | Hq-1 | 2023-05-24 07:54 | 2023-05-24 08:09 | 2023-05-24 09:55 | 2023-05-25 10:00 | Breakdown | 3312-337 Vacumn pressure timeout | Change HCB stage pump, clean vacuum container | Offline | Mr. Tú, Mr. Tới | BME | Done | Change retic to module 2 for running on fixing state | TRUE | QC: 12 (6 Hq00699, 6 Hq00692) | 12 | report_Images/80e0c740.Image.032608.jpg
2305221536 | Molecular Biology | QuantStudio 7 Pro | QS7P-02 | 2023-02-16 15:35 | 2023-02-16 15:35 | 2023-05-23 09:30 | 2023-05-24 17:35 | Breakdown | #0x8865 - Unable to power up instrument | Change Main board | Offline | Mr. Danh, Mr. Jeffry | BME | Done |  | FALSE |  |  | report_Images/2305221536.Image.083705.jpg
b0f1270a | AO | Tempus | Tempus600 | 2023-05-24 12:27 | 2023-05-24 12:28 | 2023-05-24 16:00 | 2023-05-24 17:30 | Breakdown | Tube tranfer bell can't move | Tight motor pulli | Offline | Mr. Duy | BME | Done |  | FALSE |  |  | report_Images/b0f1270a.Image.114358.jpg
906fcafd | Bio-Immuno | Alinity C | AC-1 | 2023-05-23 09:20 | 2023-05-23 09:23 | 2023-05-23 09:25 | 2023-05-23 21:45 | Breakdown | Reagent Carousel blocked | Take out reagent cartridges, re-calibration carousel | Online | Mr. Vũ | BME | Done |  | FALSE |  |  | report_Images/906fcafd.Image.124623.jpg
3dabcf55 | Bio-Immuno | Alinity I | AI-1 | 2023-05-23 12:33 | 2023-05-23 12:40 | 2023-05-23 12:43 | 2023-05-23 13:30 | Breakdown | R2 pipettor pump motor error | Change pump, flush wash | Online | Mr. Vũ | BME | Done |  | FALSE |  |  | report_Images/3dabcf55.Image.165401.jpg
2305221535 | Molecular Biology | QuantStudio 7 Pro | QS7P-01 | 2023-02-16 15:35 | 2023-02-16 15:35 | 2023-05-22 09:30 | 2023-05-22 17:35 | Breakdown | #C8148 - Cover could not be lower | Replace sensor, main board, reconnect cables | Offline | Mr. Danh, Mr. Jeffry | BME | Done |  | FALSE |  |  | Form_Images/2305221535QS7P-01.Image.084621.jpg
2305221 | All | Sample Ladder | Lift | 2023-05-22 11:55 |  | 2023-05-22 13:30 | 2023-05-22 15:00 | Breakdown | Counterweight's cable of the ladder door is broken | Connect the cable | Offline | HS | Mr. Lâm | Done |  | FALSE |  |  | Form_Images/2305221419Lift.Image.072722.jpg
2305222 | GLP | Centrifuge Module | CM-03 | 2023-05-22 13:30 |  | 2023-05-22 13:40 | 2023-05-22 14:40 | Breakdown | AP don't accept CAR | Modify AP parameters | Offline | BME | Mr. Hoàng | Done |  | FALSE
c4b739a7 | GLP | Centrifuge Module | CM-03 | 2023-05-22 13:30 | 2023-05-22 13:30 | 2023-05-22 14:12 | 2023-05-22 14:30 | Breakdown | AP can't work as expected | Adjust catch position for AP | Offline | Mr. Hoàng | BME | Done |  | FALSE |  |  | report_Images/c4b739a7.Image.040123.png
ad502970 | GLP | Aliquoter Module | AM-01 | 2023-05-22 12:55 | 2023-05-22 12:57 | 2023-05-22 12:57 | 2023-05-22 13:10 | Breakdown | Sample can't go in to Aliquoter Module | Reset ALQ module controller | Offline | Mr. Hoàng | BME | Done |  | FALSE |  |  | report_Images/ad502970.Image.040459.png
2305223 | GLP | Stramax Connect Module | STR-01 | 2023-05-22 08:00 |  | 2023-05-22 08:10 | 2023-05-22 09:40 | Breakdown | Cuvette pusher blocked | Take out all cuvette, clean area | Online | Mr. Sơn | Mr. Huy | Done |  | FALSE |  |  | Form_Images/2305221044STR-01.Image.034733.jpg
3ce348cc | GLP | Cars and Track | CTR | 2023-05-19 17:04 | 2023-05-19 17:04 | 2023-05-19 05:04 | 2023-05-19 18:14 | Maintenance | Prevent maintenance | Maintenance | Offline | BME |  | Done
2305191140 | AO | Tempus | Tempus600 | 2023-05-19 11:33 |  | 2023-05-19 11:36 | 2023-05-19 11:40 | Breakdown | Sample not send from carousel | Check carousel, re-align and re-power |  | BME |  | Done |  | FALSE |  |  | Form_Images/2305191140Tempus600.Image.044328.jpg
2305191132 | GLP | Tempus Module | IOM_TIO-02 | 2023-05-19 09:32 |  | 2023-05-19 09:43 | 2023-05-19 09:55 | Breakdown | Tube out of position, finger collision -> broken | Change finger | Offline | BME |  | Done |  | FALSE |  |  | Form_Images/2305191132IOM_TIO-02.Image.044003.jpg
2305181335 | GLP | Loader Module | LM-01 | 2023-05-18 13:10 |  | 2023-05-18 13:15 | 2023-05-18 13:35 | Breakdown | Module don't work | Re-start module |  | BME |  | Follow up | App freeze | FALSE |  |  | Form_Images/2305181335LM-01.Image.063807.jpg
2305180134 | GLP | Tempus Module | IOM_TIO-01 | 2023-05-17 17:29 |  | 2023-05-17 17:34 | 2023-05-17 17:50 | Breakdown | Change rubber | Change Break's Rubber | Offline | BME |  | Done |  | FALSE
2305180137 | GLP | Tempus Module | IOM_TIO-02 | 2023-05-17 17:29 |  | 2023-05-17 17:34 | 2023-05-17 17:50 | Breakdown | Change rubber | Change Break's Rubber | Offline | BME |  | Done |  | FALSE
13af91ee | Bio-Immuno | Thermal hygrometer | AK-05 | 2023-05-16 15:31 | 2023-05-16 15:41 | 2023-05-16 16:11 | 2023-05-16 17:11 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE |  |  | report_Images/13af91ee.Image.101207.png
429c2325 | Bio-Immuno | Thermal hygrometer | AK-04 | 2023-05-16 15:29 | 2023-05-16 15:39 | 2023-05-16 16:09 | 2023-05-16 17:09 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE |  |  | report_Images/429c2325.Image.101024.png
698426ab | Bio-Immuno | Thermal hygrometer | AK-01 | 2023-05-16 15:28 | 2023-05-16 15:38 | 2023-05-16 16:08 | 2023-05-16 17:08 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE |  |  | report_Images/698426ab.Image.100925.png
c3345119 | R&D | Thermal hygrometer | AK-03 | 2023-05-16 15:27 | 2023-05-16 15:37 | 2023-05-16 16:07 | 2023-05-16 17:07 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE |  |  | report_Images/c3345119.Image.100831.png
5aac0ef7 | Microbiology | Thermo hygrometer | NAK-02 | 2023-05-16 15:16 | 2023-05-16 15:26 | 2023-05-16 15:56 | 2023-05-16 16:56 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE |  |  | report_Images/5aac0ef7.Image.095709.png
5426986f | ELISA | Liquid in glass Thermometer | NK-14 | 2023-05-16 15:07 | 2023-05-16 15:17 | 2023-05-16 15:47 | 2023-05-16 16:47 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE |  |  | report_Images/5426986f.Image.094836.png
449ee5d4 | ELISA | Liquid in glass Thermometer | NK-10 | 2023-05-16 15:06 | 2023-05-16 15:16 | 2023-05-16 15:46 | 2023-05-16 16:46 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE |  |  | report_Images/449ee5d4.Image.094721.png
c1277417 | ELISA | Liquid in glass Thermometer | NK-12 | 2023-05-16 15:04 | 2023-05-16 15:14 | 2023-05-16 15:44 | 2023-05-16 16:44 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE |  |  | report_Images/c1277417.Image.094536.png
56b734a8 | Microbiology | Thermometer | NK - 03 | 2023-05-16 15:03 | 2023-05-16 15:13 | 2023-05-16 15:43 | 2023-05-16 16:43 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE |  |  | report_Images/56b734a8.Image.094437.png
29051186 | Microbiology | Thermometer | NK - 02 | 2023-05-16 15:02 | 2023-05-16 15:12 | 2023-05-16 15:42 | 2023-05-16 16:42 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE |  |  | report_Images/29051186.Image.094319.png
cccf4790 | Hematology - Coagulation | Timer | TI-02 | 2023-05-16 14:58 | 2023-05-16 15:08 | 2023-05-16 15:38 | 2023-05-16 16:38 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE |  |  | report_Images/cccf4790.Image.093925.png
079211ad | R&D | Pipette 1-10ul | F2-PP10 | 2023-05-16 14:37 | 2023-05-16 14:47 | 2023-05-16 15:17 | 2023-05-16 16:17 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE |  |  | report_Images/079211ad.Image.091829.png
ee93729b | R&D | Pipette 100-1000ul | F2-PP1000 | 2023-05-16 14:36 | 2023-05-16 14:46 | 2023-05-16 15:16 | 2023-05-16 16:16 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE |  |  | report_Images/ee93729b.Image.091744.png
e6c3481e | ELISA | Pipette proline | PP-12 | 2023-05-16 14:36 | 2023-05-16 14:46 | 2023-05-16 15:16 | 2023-05-16 16:16 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE |  |  | report_Images/e6c3481e.Image.091652.png
543626b7 | ELISA | Finnpippet 500 microlite | PP-006 | 2023-05-16 14:24 | 2023-05-16 14:34 | 2023-05-16 15:04 | 2023-05-16 16:04 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE |  |  | report_Images/543626b7.Image.090523.png
cde9b173 | R&D | Pipette 20-200ul | F2-PP200 | 2023-05-16 14:23 | 2023-05-16 14:33 | 2023-05-16 15:03 | 2023-05-16 16:03 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE |  |  | report_Images/cde9b173.Image.090428.png
7de80c3b | ELISA | Finnpipette 50 microlite | PP-002 | 2023-05-16 14:22 | 2023-05-16 14:32 | 2023-05-16 15:02 | 2023-05-16 16:02 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE |  |  | report_Images/7de80c3b.Image.090259.png
57ac026e | Microbiology | Pipette 10-100 | PP - 02 | 2023-05-16 14:21 | 2023-05-16 14:31 | 2023-05-16 15:01 | 2023-05-16 16:01 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE |  |  | report_Images/57ac026e.Image.090202.png
54008da5 | ELISA | Pippette | PP-008 | 2023-05-16 14:20 | 2023-05-16 14:30 | 2023-05-16 15:00 | 2023-05-16 16:00 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE |  |  | report_Images/54008da5.Image.090053.png
7bcf5a92 | Microbiology | Pipette 0.5-10 | PP - 03 | 2023-05-16 14:13 | 2023-05-16 14:23 | 2023-05-16 14:53 | 2023-05-16 15:53 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE |  |  | report_Images/7bcf5a92.Image.085937.png
cd3eb1a6 | Molecular Biology | Pipette 10-100uL | PP-02 | 2023-05-16 14:03 | 2023-05-16 14:13 | 2023-05-16 14:43 | 2023-05-16 15:43 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE |  |  | report_Images/cd3eb1a6.Image.084405.png
9ef5a45c | ELISA | Humanpette 50 microlite | PP- 04 | 2023-05-16 13:59 | 2023-05-16 14:09 | 2023-05-16 14:39 | 2023-05-16 15:39 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE |  |  | report_Images/9ef5a45c.Image.084104.png
9ef5a45d | Molecular Biology | Multichannel Pipette (5-100uL) | PP-04 | 2023-05-16 14:00 | 2023-05-16 14:09 | 2023-05-16 14:39 | 2023-05-16 15:39 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE |  |  | report_Images/9ef5a45d.Image.084836.png
c5ab549a | Molecular Biology | Multichannel Pipette (5-50uL) | PP-05 | 2023-05-16 13:20 | 2023-05-16 13:30 | 2023-05-16 14:00 | 2023-05-16 15:00 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE |  |  | report_Images/c5ab549a.Image.080040.png
a692857b | Molecular Biology | Multichannel Pipette (30-300uL) | PP-03 | 2023-05-16 13:18 | 2023-05-16 13:28 | 2023-05-16 13:58 | 2023-05-16 14:58 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE |  |  | report_Images/a692857b.Image.075931.png
2305172114 | Bio-Immuno | Alinity I | AI-3 | 2023-05-16 09:00 |  | 2023-05-16 09:09 | 2023-05-16 09:50 | Breakdown | Shutter sensor error | Clean shutter, sensor; Test shutter open/ close | Offline | Mr. Vũ | BME | Done |  | FALSE |  |  | Form_Images/2305172114AI-3.Image.141828.png
2305172122 | Bio-Immuno | Alinity I | AI-2 | 2023-05-16 06:20 |  | 2023-05-16 06:45 | 2023-05-16 07:10 | Breakdown | Sample prove out of alignment | Replace sample probe | Offline | Mr. Vũ | BME | Done |  | FALSE |  |  | Form_Images/2305172122AI-2.Image.142358.png
2305131552 | Bio-Immuno | Cobas e801 Analytical unit | MD-09 | 2023-05-13 14:37 |  | 2023-05-13 14:50 | 2023-05-13 15:30 | Breakdown | Cup sensor error | Check and clean sensor and around area, check performance | Online | Mr. Khang | BME | Done |  | FALSE |  |  | Form_Images/2305131552MD-09.Image.085735.jpg
2305151002 | GLP | Cars and Track | CTR | 2023-05-12 13:55 |  | 2023-05-12 14:00 | 2023-05-12 15:00 | Maintenance | Prevent Maintenance | Maintenance |  | BME |  | Done |  | FALSE
2305121339 | GLP | Recapper Module | RM-01 | 2023-05-12 13:34 |  | 2023-05-12 13:39 | 2023-05-12 14:00 | Breakdown | Robotic failed | Clean and realignt | Offline | Mr. Hoàng | BME | Done |  | FALSE |  |  | Form_Images/2305121339RM-01.Image.064040.jpg
2305172118 | Bio-Immuno | Alinity C | AC-3 | 2023-05-11 14:44 |  | 2023-05-11 14:45 | 2023-05-11 15:00 | Breakdown | R2 pipettor movement restricted | Replace R2 probe | Online | Mr. Vũ | BME | Done |  | FALSE |  |  | Form_Images/2305172118AC-3.Image.142123.png
2305111016 | AO | Tempus | Tempus600 | 2023-05-11 09:00 |  | 2023-05-11 09:05 | 2023-05-11 10:35 | Breakdown | Tube dropped | Check and re-process | Offline | Mr. Hoang | BME | Done |  | FALSE |  |  | Form_Images/2305111016Tempus600.Image.032119.jpg
6ec0f42d | ELISA | DAR800 | ELISA-02 | 2023-05-10 12:28 | 2023-05-10 12:38 | 2023-05-10 13:08 | 2023-05-10 14:08 | Calibration | Annual Calibration | Calibration | Offline | EDC-HCM | BME | Done |  | FALSE |  |  | report_Images/6ec0f42d.Image.071016.png
2305100914 | GLP | Cars and Track | CTR | 2023-05-10 12:00 |  | 2023-05-10 12:20 | 2023-05-10 12:22 | Breakdown | Lots of car stuck | Change track parameters |  | BME |  | Done |  | FALSE |  |  | Form_Images/2305100914CTR.Image.021721.jpg
2305100911 | GLP | Input/ Output Module | IOM-03 | 2023-05-10 09:06 |  | 2023-05-10 09:11 | 2023-05-10 09:15 | Breakdown | Display freeze | Reboot machine |  | BME |  | Done |  | FALSE |  |  | Form_Images/2305100911IOM-03.Image.021400.jpg
2305091535 | GLP | Storage on SM | STORE-01 | 2023-05-09 15:30 |  | 2023-05-09 15:35 | 2023-05-09 15:52 | Breakdown | Left door stuck | Clean rail and restart |  | BME |  | Done |  | FALSE |  |  | Form_Images/2305091535STORE-01.Image.083704.jpg
2305091252 | GLP | Tempus Module | IOM_TIO-01 | 2023-05-09 12:47 |  | 2023-05-09 12:52 | 2023-05-09 13:15 | Breakdown | Change rubber | Change Break's Rubber | Offline | BME |  | Done |  | FALSE |  |  | Form_Images/2305091252IOM_TIO-01.Image.055309.jpg
2305091253 | GLP | Tempus Module | IOM_TIO-02 | 2023-05-09 12:48 |  | 2023-05-09 12:53 | 2023-05-09 13:15 | Breakdown | Change rubber | Change Break's Rubber | Offline | BME |  | Done |  | FALSE |  |  | Form_Images/2305091253IOM_TIO-02.Image.055348.jpg
2305171632 | Bio-Immuno | Immulite | MD-08 | 2023-05-05 08:30 |  | 2023-05-05 09:00 | 2023-05-05 09:32 | Breakdown | System cannot startup | Check main board, UPS -> out of power | Offline | Mr. Thụân An | Ms. Huyền | Done |  | FALSE |  |  | Form_Images/2305171632MD-08.Image.093629.jpg
2305041815 | Molecular Biology | Momentum | MOM | 2023-05-04 15:55 |  | 2023-05-04 16:00 | 2023-05-04 18:00 | Breakdown | New parameters checking | Check and re-teach mover point |  | BME |  | Done |  | FALSE
2305041337 | GLP | Cars and Track | CTR | 2023-05-04 13:32 |  | 2023-05-04 13:37 | 2023-05-04 14:37 | Maintenance | Prevent Maintenance | Maintenance |  | BME |  | Done |  | FALSE
3515 | Hematology - Coagulation | Microcope NIKON | KHV-02 | 2023-05-04 13:55 | 2023-05-04 13:55 | 2023-05-04 13:55 | 2023-05-04 14:35 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Kha | BME | Done
15325 | Microbiology | Microscope | KHV-03 | 2023-05-04 13:55 | 2023-05-04 13:55 | 2023-05-04 13:55 | 2023-05-04 14:35 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Kha | BME | Done
2305041053 | Bio-Immuno | Alinity C | AC-3 | 2023-05-04 10:50 |  | 2023-05-04 10:53 | 2023-05-04 11:20 | Breakdown | App freeze | Wait and restart PC | Online | Mr. Vũ | BME | Done |  | TRUE |  |  | Form_Images/2305041053AC-3.Image.035721.jpg
2305041059 | GLP | Storage on SM | STORE-01 | 2023-05-04 10:54 |  | 2023-05-04 10:59 | 2023-05-04 11:10 | Breakdown | Device don't work | Restart | Online | Mr. Hoàng | BME | Done |  | FALSE |  |  | Form_Images/2305041059STORE-01.Image.040100.jpg
2305031612 | Bio-Immuno | RO system Abbott | RO-HP | 2023-05-03 13:55 |  | 2023-05-03 14:00 | 2023-05-03 15:00 | Breakdown | Water leaked | Fix tight, change filtering coarse | Offline | Mr. Tâm | Mrs. Tuyên | Done |  | FALSE |  |  | Form_Images/2305031612RO-HP.Image.091521.jpg
2305031605 | Bio-Immuno | Alinity C | AC-3 | 2023-05-03 10:30 |  | 2023-05-03 11:00 | 2023-05-03 11:20 | Breakdown | Sample probe bended | Change probe | Offline | Mr. Tới | Ms. Trân | Done | Decap failed - lot 2244015-2023-12-31 | FALSE |  |  | Form_Images/2305031605AC-3.Image.091145.jpg
2305171050 | Microbiology | Class II Biosafety Cabinet | ATSH-02 | 2023-04-28 15:39 | 2023-04-28 15:49 | 2023-04-28 16:19 | 2023-04-28 23:30 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE |  |  | Form_Images/2305171050ATSH-02.Image.035756.png
2305171100 | Microbiology | Class II Biosafety Cabinet | ATSH-03 | 2023-04-28 15:39 | 2023-04-28 15:49 | 2023-04-28 16:19 | 2023-04-28 23:30 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE |  |  | Form_Images/2305171100ATSH-03.Image.040108.png
2305171051 | Microbiology | Class II Biosafety Cabinet | ATSH-04 | 2023-04-28 15:39 | 2023-04-28 15:49 | 2023-04-28 16:19 | 2023-04-28 23:30 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE |  |  | Form_Images/2305171051ATSH-04.Image.035810.png
2305171058 | Molecular Biology | Class II Biosafety Cabinet | BSC-01 | 2023-04-28 15:39 | 2023-04-28 15:49 | 2023-04-28 16:19 | 2023-04-28 23:30 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE |  |  | Form_Images/2305171058BSC-01.Image.035938.png
2305171101 | Microbiology | Electronic scales | CAN-01 | 2023-04-28 15:39 | 2023-04-28 15:49 | 2023-04-28 16:19 | 2023-04-28 23:30 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE |  |  | Form_Images/2305171101CAN-01.Image.040148.png
2305171046 | Cyto Lab | FUME HOOD | FH-01 | 2023-04-28 15:39 | 2023-04-28 15:49 | 2023-04-28 16:19 | 2023-04-28 23:30 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE |  |  | Form_Images/2305171046FH-01.Image.035745.png
25399c03 | Cyto Lab | OVEN | OVEN-01 | 2023-04-28 15:39 | 2023-04-28 15:49 | 2023-04-28 16:19 | 2023-04-28 17:19 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE |  |  | report_Images/25399c03.Image.102023.png
5906258c | Bio-Immuno | Refrigerator HAIER | TL-011 | 2023-04-28 15:37 | 2023-04-28 15:47 | 2023-04-28 16:17 | 2023-04-28 17:17 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE |  |  | report_Images/5906258c.Image.101837.png
671b73fa | Warehouse | Refrigerator HAIER | TL-012 | 2023-04-28 15:36 | 2023-04-28 15:46 | 2023-04-28 16:16 | 2023-04-28 17:16 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE |  |  | report_Images/671b73fa.Image.101707.png
2305181252 | Molecular Biology | Centrifuge | LT-01 | 2023-04-28 13:55 | 2023-04-28 14:00 | 2023-04-28 14:00 | 2023-04-28 16:30 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE |  |  | Form_Images/2305181252LT-01.Image.055309.png
2305181245 | AO | Centrifuge | LT-011 | 2023-04-28 13:55 | 2023-04-28 14:00 | 2023-04-28 14:00 | 2023-04-28 16:30 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE |  |  | Form_Images/2305181245LT-011.Image.054701.png
2305181242 | AO | Centrifuge | LT-012 | 2023-04-28 13:55 | 2023-04-28 14:00 | 2023-04-28 14:00 | 2023-04-28 16:30 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE |  |  | Form_Images/2305181242LT-012.Image.054517.png
2305181253 | Bio-Immuno | Cold centrifuge | LTL-01 | 2023-04-28 13:55 | 2023-04-28 14:00 | 2023-04-28 14:00 | 2023-04-28 16:30 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE |  |  | Form_Images/2305181253LTL-01.Image.055510.png
2305181256 | Bio-Immuno | VORTEX machine | ML- 01 | 2023-04-28 13:55 | 2023-04-28 14:00 | 2023-04-28 14:00 | 2023-04-28 16:30 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE |  |  | report_Images/2305181256.Image.092725.png
2305181257 | ELISA | VORTEX machine | ML-02 | 2023-04-28 13:55 | 2023-04-28 14:00 | 2023-04-28 14:00 | 2023-04-28 16:30 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE |  |  | Form_Images/2305181257ML-02.Image.055743.png
2305181255 | R&D | Máy lắc Microplate | MP01 | 2023-04-28 13:55 | 2023-04-28 14:00 | 2023-04-28 14:00 | 2023-04-28 16:30 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE |  |  | report_Images/2305181255.Image.092040.png
123213 | Microbiology | Autoclaves SA - 300VF | NH-01 | 2023-04-28 13:55 | 2023-04-28 14:00 | 2023-04-28 14:00 | 2023-04-28 16:30 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE
2305181259 | Microbiology | Autoclaves SA - 300VF | NH-02 | 2023-04-28 13:55 | 2023-04-28 14:00 | 2023-04-28 14:00 | 2023-04-28 16:30 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE |  |  | Form_Images/2305181259NH-02.Image.055940.png
2305181258 | Microbiology | Loop sterilizer | NH-03 | 2023-04-28 13:55 | 2023-04-28 14:00 | 2023-04-28 14:00 | 2023-04-28 16:30 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE |  |  | Form_Images/2305181258NH-03.Image.055844.png
74091334 | Microbiology | Electronic scales | CAN-01 | 2023-04-28 13:12 | 2023-04-28 14:00 | 2023-04-28 13:52 | 2023-04-28 14:52 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE |  |  | report_Images/74091334.Image.075252.png
bbe9a311 | ELISA | Refrigerator SANAKY | TL-007 | 2023-04-28 13:10 | 2023-04-28 14:00 | 2023-04-28 13:50 | 2023-04-28 14:50 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE |  |  | report_Images/bbe9a311.Image.075138.png
5b198520 | Microbiology | Refrigerator Sanaky | TL-01 | 2023-04-28 13:09 | 2023-04-28 14:00 | 2023-04-28 13:49 | 2023-04-28 14:49 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE |  |  | report_Images/5b198520.Image.075025.png
707c5c9d | Hematology - Coagulation | Refrigerator SANAKY 800L | TL-008 | 2023-04-28 13:08 | 2023-04-28 14:00 | 2023-04-28 13:48 | 2023-04-28 14:48 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE |  |  | report_Images/707c5c9d.Image.074857.png
c8bc7016 | Bio-Immuno | Refrigerator Sanyo | TL-002 | 2023-04-28 13:07 | 2023-04-28 13:17 | 2023-04-28 13:47 | 2023-04-28 14:47 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE |  |  | report_Images/c8bc7016.Image.074743.png
9a40b371 | Microbiology | Refrigerator Sanaky | TL-02 | 2023-04-28 13:05 | 2023-04-28 13:15 | 2023-04-28 13:45 | 2023-04-28 14:45 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE |  |  | report_Images/9a40b371.Image.074628.png
dc915ae7 | Bio-Immuno | Refrigerator PANASONIC | TL-010 | 2023-04-28 13:04 | 2023-04-28 13:14 | 2023-04-28 13:44 | 2023-04-28 14:44 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE |  |  | report_Images/dc915ae7.Image.074506.png
e3e91680 | Microbiology | Firlabo Incubator 37°C | TU-01 | 2023-04-28 13:02 | 2023-04-28 13:12 | 2023-04-28 13:42 | 2023-04-28 14:42 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE |  |  | report_Images/e3e91680.Image.074239.png
b948cf13 | ELISA | INCUBATOR | TU-02 | 2023-04-28 13:02 | 2023-04-28 13:12 | 2023-04-28 13:42 | 2023-04-28 14:42 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE |  |  | report_Images/b948cf13.Image.074338.png
2304271534 | Hematology - Coagulation | Minicap Flex Piercing | DD-03 | 2023-04-27 11:29 |  | 2023-04-27 12:05 | 2023-04-27 15:42 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Huy | BME | Done | Probe 1 still can't work | FALSE
2304271510 | Microbiology | Urine Cell Analyzer UF-4000 | UF-01 | 2023-04-27 10:51 |  | 2023-04-27 13:22 | 2023-04-27 13:58 | Breakdown | UPS out of batteries | Replace the batteries for UPS | Offline | Mr. Đại | Ms. Hân | Done |  | FALSE
2304271827 | GLP | Centrifuge Module | CM-03 | 2023-04-27 12:55 |  | 2023-04-27 13:00 | 2023-04-27 13:15 | Breakdown | AP can't access CAR | Modify parameters | Offline | Mr. Hoàng | BME | Done |  | FALSE
2304261926 | Hematology - Coagulation | Rida qLine Allergy | Rida | 2023-04-26 14:55 |  | 2023-04-26 15:00 | 2023-04-26 15:30 | Maintenance | Prevent maintenance | Maintenance | Offline | Mr. Quân | BME | Done | Scheduled | FALSE
2304271226 | Hematology - Coagulation | Rida qLine Allergy | Rida | 2023-04-26 13:55 |  | 2023-04-26 14:00 | 2023-04-26 14:30 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Quân | BME | Done | Scheduled | FALSE
2304261832 | Hematology - Coagulation | Tosoh HLC-723 G11 LA analyzer | HbA1c-02 | 2023-04-26 12:27 |  | 2023-04-26 12:32 | 2023-04-26 13:12 | Breakdown | Pipe leaking, drag sensor error | Change pipe and sensor | Offline | Mr. Đạt | BME | Done | Done from 21/4 | FALSE |  |  | Form_Images/2304261832Tosoh HLC-723 G11 LA anal_14189497602842954601.Image.113414.jpg
2304261846 | GLP | Centrifuge Module | CM-03 | 2023-04-26 10:55 |  | 2023-04-26 11:00 | 2023-04-26 11:10 | Breakdown | Car stop on AP | Take out and manual running samples, align and re initialize AP |  | BME |  | Done |  | FALSE
2304261844 | Molecular Biology | Grenova Washer | GNV-01 | 2023-04-25 12:55 |  | 2023-04-25 13:00 | 2023-04-25 13:20 | Breakdown | Can't accept SD | Using instrument protocol |  | BME |  | Done | Watch for any miss operation | FALSE
2304261830 | Hematology - Coagulation | Minicap Flex Piercing | DD-03 | 2023-04-24 08:54 |  | 2023-04-24 08:59 | 2023-04-24 16:46 | Breakdown | Continute wash electrophosis head washing | Wash electrophosis head | Offline | Mr. Huy | BME | Done | One electrophosis head still can't work, Using only one electrophosis head | FALSE |  |  | Form_Images/2304261830Minicap Flex Piercing- DD-03.Image.113421.jpg
2304261841 | GLP | Cars and Track | CTR | 2023-04-24 13:35 |  | 2023-04-24 13:40 | 2023-04-24 14:40 | Maintenance | Prevent Maintenance | Maintenance |  | BME |  | Done |  | FALSE
2304271551 | Hematology - Coagulation | Tosoh HLC-723 G11 LA analyzer | HbA1c-02 | 2023-04-24 08:00 |  | 2023-04-24 08:20 | 2023-04-24 12:00 | Breakdown | Can not load software; drain full error | Change main board fuse, change pump | Offline | Mr. Đạt | BME | Done | QC pass | FALSE
2305180236 | AO | Tempus | Tempus600 | 2023-04-24 09:36 |  | 2023-04-24 10:00 | 2023-04-24 12:00 | Breakdown | Tube droped | Tight transfer bell, clean transport pipes | Offline | Mr. Hoàng | BME | Done |  | FALSE |  |  | Form_Images/2305180236Tempus600.Image.194120.png
2304261842 | Hematology - Coagulation | Minicap Flex Piercing | DD-03 | 2023-04-20 16:32 |  | 2023-04-21 09:18 | 2023-04-21 16:40 | Breakdown | Pressure error | Change valve, wash 1 electrophosis head that stuck | Offline | Mr. Huy | BME | Done | Valve changed; One electrophosis head can't work; stop from 8/3/2023 | FALSE
135 | GLP | Loader Module | LM-01 | 2023-04-20 16:55 |  | 2023-04-20 17:00 | 2023-04-20 17:30 | Breakdown | Can not initialize Storage Left | Open Storage, check sensor and clean, re-initialize |  | BME |  | Done |  | FALSE
133 | Bio-Immuno | Alinity I | AI-3 | 2023-04-20 13:55 |  | 2023-04-20 12:00 | 2023-04-20 12:15 | Breakdown | R1 probe can't aspirate | Change R1 Probe | Online | Mr. Vũ | BME | Done |  | FALSE
6 | Bio-Immuno | Alinity C | AC-3 | 2023-04-19 12:55 |  | 2023-04-19 13:00 | 2023-04-19 16:00 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Vũ | BME | Done | Scheduled | FALSE
132 | Hematology - Coagulation | Tosoh HLC-723 G11 LA analyzer | HbA1c-02 | 2023-04-19 12:00 |  | 2023-04-19 12:20 | 2023-04-19 14:00 | Breakdown | Baseline does not stability -> Valve 4 pipe break | fix pipe | Offline | Mr. Bình | BME | Done | Wait to have new pipe to change; Replace on 26/4/2023 | TRUE | HbA1c | 7
116 | AO | Tempus | Tempus600 | 2023-04-13 17:30 |  | 2023-04-13 18:00 | 2023-04-13 19:00 | Breakdown | Dust on the transfer pipe | Clean pipe, change rubbers | Offline | Mr. Hoàng, Mr. Duy | BME | Follow up |  | FALSE
114 | Bio-Immuno | Alinity C | AC-3 | 2023-04-13 14:55 |  | 2023-04-13 15:00 | 2023-04-13 15:15 | Breakdown | Change Dry tip | Change Dry tip |  | BME |  | Done |  | FALSE
112 | Bio-Immuno | Alinity C | AC-1 | 2023-04-13 14:55 |  | 2023-04-13 15:00 | 2023-04-13 15:10 | Breakdown | Change Dry tip | Change Dry tip |  | BME |  | Done |  | FALSE
113 | Bio-Immuno | Alinity C | AC-2 | 2023-04-13 14:55 |  | 2023-04-13 15:00 | 2023-04-13 15:10 | Breakdown | Change Dry tip | Change Dry tip |  | BME |  | Done |  | FALSE
2304271548 | Hematology - Coagulation | Alinity H-series | Hq-1 | 2023-04-13 09:00 |  | 2023-04-13 09:20 | 2023-04-13 11:00 | Breakdown | Extended autoclean failure | Clean waste pipe | Offline | Mr. Tới | BME | Done | Run Extended autoclean pass | FALSE
107 | GLP | Cars and Track | CTR | 2023-04-12 10:55 |  | 2023-04-12 11:00 | 2023-04-12 12:00 | Maintenance | Prevent Maintenance | Maintenance |  | BME |  | Done |  | FALSE
111 | AO | Tempus | Tempus600 | 2023-04-11 17:00 |  | 2023-04-11 18:00 | 2023-04-11 19:00 | Breakdown | Tube out of position -> miss tube | Align tube transfer cover on transfer line | Offline | Mr. Hoàng | BME | Done | 1 sample tube out missed -> founded | FALSE
2304271545 | Hematology - Coagulation | Alinity H-series | Hq-1 | 2023-04-11 09:00 |  | 2023-04-11 09:30 | 2023-04-11 11:30 | Breakdown | Purge vaccum buffer tank issue | Clean Vaccum tube, test vaccum | Offline | Mr. Tới | BME | Done |  | FALSE
2304271542 | Hematology - Coagulation | Alinity H-series | Hq-1 | 2023-04-09 10:00 |  | 2023-04-09 10:20 | 2023-04-09 11:00 | Breakdown | Vaccum buffer tank issue | Check and clean tube and valve, Fit reseating cables | Offline | Mr. Tuấn Anh | BME | Done |  | FALSE
110 | GLP | Tempus Module | IOM_TIO-02 | 2023-04-08 09:55 |  | 2023-04-08 10:00 | 2023-04-08 10:15 | Breakdown | Finger stuck with tube | Take out tube, clean and change Finger | Offline | BME |  | Done |  | FALSE
105 | GLP | Tempus Module | IOM_TIO-01 | 2023-04-07 16:55 |  | 2023-04-07 17:00 | 2023-04-07 17:30 | Breakdown | Change rubber | Change Break's Rubber | Offline | BME |  | Done |  | FALSE
106 | GLP | Tempus Module | IOM_TIO-02 | 2023-04-07 16:55 |  | 2023-04-07 17:00 | 2023-04-07 17:20 | Breakdown | Change rubber | Change Break's Rubber | Offline | BME |  | Done |  | FALSE
8151 | Bio-Immuno | Cobas e801 Analytical unit | MD-09 | 2023-04-07 14:55 |  | 2023-04-07 15:00 | 2023-04-07 15:20 | Breakdown | Connect server | Config server | Online | Mr. Khang | BME | Done |  | FALSE
100 | Bio-Immuno | Alinity I | AI-4 | 2023-04-07 13:55 |  | 2023-04-07 14:00 | 2023-04-07 14:20 | Breakdown | Reagent moving stuck | Check movement, reinitialize module | Online | Mr. Vũ | BME | Done |  | FALSE
102 | Molecular Biology | ABI 7500 | ABI-04 | 2023-04-07 10:55 |  | 2023-04-07 11:00 | 2023-04-07 11:10 | Breakdown | Adapter crack | Still using | Online | SISC team | BME | Done | Have quotation for change but still can use | TRUE | Sars
103 | GLP | Cars and Track | CTR | 2023-04-06 15:55 |  | 2023-04-06 16:00 | 2023-04-06 17:00 | Breakdown | A lot of stuck | Take out fail Car and fix |  | BME |  | Done |  | FALSE
122 | Cyto Lab | FUME HOOD | FH-01 | 2023-04-04 09:55 |  | 2023-04-04 10:00 | 2023-04-06 10:00 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Thắng | Ms. Hà | Done | Scheduled | FALSE
115 | GLP | Recapper Module | RM-01 | 2023-04-05 12:55 |  | 2023-04-05 13:00 | 2023-04-05 15:00 | Breakdown | AP can't move correctly, clutch plate broken | Clutch plate aluminium | Offline | Mr. Hoàng | BME | Done |  | FALSE
104 | GLP | Recapper Module | RM-01 | 2023-04-04 19:55 |  | 2023-04-04 20:00 | 2023-04-04 20:30 | Breakdown | AP motion fail --> crack Car's pin | Stop using 1 module | Online | BME | Mr. Viễn | Done | Need to fix AP motor connecter | FALSE
98 | Molecular Biology | Momentum | MOM | 2023-04-04 10:55 |  | 2023-04-04 11:00 | 2023-04-04 14:00 | Breakdown | Create new script for running | Modify and sorting for data processing on software |  | BME |  | Done |  | FALSE
99 | GLP | Cars and Track | CTR | 2023-04-04 12:55 |  | 2023-04-04 13:00 | 2023-04-04 14:00 | Breakdown | CAR not enough on processing module | Increase CAR limit |  | BME |  | Done | on TWM | FALSE
97 | Molecular Biology | QuantStudio 7 Pro | QS7P-01 | 2023-04-04 10:55 |  | 2023-04-04 11:00 | 2023-04-04 11:20 | Breakdown | Cover could not be lowered | Change block, check around of block |  | BME |  | Follow up | Wait to repair | FALSE
123 | Cyto Lab | OVEN | OVEN-01 | 2023-04-04 09:55 |  | 2023-04-04 10:00 | 2023-04-04 10:30 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Thắng | Ms. Hà | Done | Scheduled | FALSE
87 | GLP | Cars and Track | CTR | 2023-04-03 13:55 |  | 2023-04-03 14:00 | 2023-04-03 15:00 | Maintenance | Prevent Maintenance | Maintenance |  | BME |  | Done |  | FALSE
93 | Hematology - Coagulation | QSmart | QSmart | 2023-04-03 08:55 |  | 2023-04-03 09:00 | 2023-04-03 11:00 | Breakdown | Install and run demo | Install and run demo | Offline | Mr. Hiền | Mr. Sơn | Done | Still not connect to LIS, Have result, need to verify | FALSE
95 | Molecular Biology | QuantStudio 7 Pro | QS7P-01 | 2023-04-03 09:55 |  | 2023-04-03 10:00 | 2023-04-03 10:20 | Breakdown | Cover count not lower | restart machine | Online | Mr. Danh | Mr. Thư | Follow up | After 4 times | FALSE
96 | Molecular Biology | QuantStudio 7 Pro | QS7P-02 | 2023-04-03 08:55 |  | 2023-04-03 09:00 | 2023-04-03 09:15 | Breakdown | Unable to power up | Repower machine |  | BME |  | Follow up | Wait to repair from Thermo | FALSE
94 | Molecular Biology | Agilent Microplate Centrifuge | MC-01 | 2023-04-03 07:55 |  | 2023-04-03 08:00 | 2023-04-03 08:20 | Breakdown | Can't connect and initially | Check connecting, teach point and confirm |  | BME |  | Done |  | FALSE
355 | Molecular Biology | KingFisher Flex | KFF-01 | 2023-03-31 15:55 |  | 2023-03-31 16:00 | 2023-03-31 17:00 | Breakdown | The heater can't move; Head movement have noise | Checking | Offline | Mr. Trí | BME | Follow up | Need to replace Heater motor, Head motor | FALSE
124 | Cyto Lab | MICROSCOPY | KHV-01 | 2023-03-31 13:55 |  | 2023-03-31 14:00 | 2023-03-31 14:30 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Hòa | Ms. Hòa | Done | Scheduled | FALSE
91 | Molecular Biology | KingFisher Flex | KFF-01 | 2023-03-30 12:55 |  | 2023-03-30 13:00 | 2023-03-30 19:30 | Breakdown | The heater can't move | Fix and clean Heater motor | Offline | Mr. Trí | BME | Follow up | Need to prepare Heater motor and Head motor for backup | FALSE
89 | ELISA | DAR800 | ELISA-02 | 2023-03-30 13:55 |  | 2023-03-30 14:00 | 2023-03-30 15:30 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Vinh | BME | Done | Scheduled | FALSE
88 | ELISA | DAW50 | ELISA-01 | 2023-03-30 13:55 |  | 2023-03-30 14:00 | 2023-03-30 15:00 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Vinh | BME | Done | Scheduled | FALSE
90 | ELISA | DAW50TS | ELISA-03 | 2023-03-30 13:55 |  | 2023-03-30 14:00 | 2023-03-30 15:00 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Vinh | BME | Done | Scheduled | FALSE
119 | Microbiology | HeliFan Plus C13 | HP-01 | 2023-03-30 12:55 |  | 2023-03-30 13:00 | 2023-03-30 14:00 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Vinh | BME | Done | Scheduled | FALSE
2304271546 | Hematology - Coagulation | Alinity H-series | Hq-2 | 2023-03-29 22:00 |  | 2023-03-29 22:20 | 2023-03-30 12:00 | Breakdown | Fail calib | Re-calib | Offline | Mr. Tới | BME | Done | QC pass | FALSE
120 | Hematology - Coagulation | Alinity H-series | Hq-1 | 2023-03-29 21:55 |  | 2023-03-29 22:00 | 2023-03-30 | Maintenance | Prevent Maintenance | Maintenance, update firmware | Offline | Mr. Tới | BME | Done | Scheduled | FALSE
121 | Hematology - Coagulation | Alinity H-series | Hq-2 | 2023-03-29 21:55 |  | 2023-03-29 22:00 | 2023-03-30 | Maintenance | Prevent Maintenance | Maintenance, update firmware | Offline | Mr. Tới | BME | Done | Scheduled | FALSE
2305171617 | Bio-Immuno | Immulite | MD-08 | 2023-03-29 12:30 |  | 2023-03-29 13:00 | 2023-03-29 17:00 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Thụân An | BME | Done | Scheduled | TRUE | AFG: 2, HCG total: 4, Estriol Unconjugated: 2, b-HCG free: 1 | 9 | Form_Images/2305171617MD-08.Image.092103.jpg
86 | Microbiology | Urine Cell Analyzer UF-4000 | UF-01 | 2023-03-29 08:00 |  | 2023-03-29 12:47 | 2023-03-29 13:55 | Breakdown | Result fail | Change Solenoid valve 35/36 | Offline | Mr. Sơn | Mr. Nguyên | Done | Done on 31/3/2023 2 pm | FALSE
125 | Cyto Lab | PATHTEZT PROCESSOR | PT-01 | 2023-03-28 07:55 |  | 2023-03-28 08:00 | 2023-03-28 18:00 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Triết | Ms. Hà | Done | Scheduled | FALSE
70 | GLP | Tempus Module | IOM_TIO-01 | 2023-03-28 16:55 |  | 2023-03-28 17:00 | 2023-03-28 17:30 | Breakdown | Change rubber | Change Break's Rubber | Offline | BME |  | Done |  | FALSE
71 | GLP | Tempus Module | IOM_TIO-02 | 2023-03-28 16:55 |  | 2023-03-28 17:00 | 2023-03-28 17:20 | Breakdown | Change rubber | Change Break's Rubber | Offline | BME |  | Done |  | FALSE
73 | GLP | Recapper Module | RM-01 | 2023-03-28 13:55 |  | 2023-03-28 14:00 | 2023-03-28 14:30 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Hoàng | BME | Done | Scheduled | FALSE
68 | Molecular Biology | Multichannel Pipette (5-50uL) | PP-05 | 2023-03-28 12:55 |  | 2023-03-28 13:00 | 2023-03-28 13:30 | Breakdown | Pippette broken | Fixed |  | BME |  | Done |  | FALSE
67 | GLP | Recapper Module | RM-01 | 2023-03-28 10:55 |  | 2023-03-28 11:00 | 2023-03-28 12:20 | Breakdown | Access point fail | Clean and realign Access Point | Offline | Mr. Hoàng | BME | Done |  | FALSE
62 | GLP | Storage on SM | STORE-01 | 2023-03-27 11:55 |  | 2023-03-27 12:00 | 2023-03-27 12:30 | Breakdown | Software not response | Checkin around and reboot | Online | Mr. Hoàng | BME | Done | Avoid repeat | FALSE
65 | GLP | Tube Assecssment Module | TAM | 2023-03-27 09:55 |  | 2023-03-27 10:00 | 2023-03-27 10:20 | Breakdown | Software not response | Check around, check log, resolve solution |  | BME |  | Done |  | FALSE
63 | GLP | Centrifuge Module | CM-02 | 2023-03-27 09:55 |  | 2023-03-27 10:00 | 2023-03-27 10:10 | Breakdown | Griper failed | Check around, take out all samples and put to GLP again to system; Clean and reboot |  | BME |  | Done |  | FALSE
64 | GLP | Tosho Connect Module | TOS-01 | 2023-03-27 08:55 |  | 2023-03-27 09:00 | 2023-03-27 09:20 | Breakdown | Connecting to LIS failed | Check around, check connection with IT team, check Moxa connect and reconect | Offline | Mr. Thành | BME | Done | Avoid power fail | FALSE
432 | GLP | Storage on SM | STORE-01 | 2023-03-24 16:55 |  | 2023-03-24 17:00 | 2023-03-24 23:30 | Breakdown | Tube error on GLP | Clean data of Kima tube to avoid Tube error on GLP | Offline | Mr. Minh | BME | Done | Clean all data and push back; finish at 11:30 pm | FALSE
58 | GLP | Cars and Track | CTR | 2023-03-24 14:55 |  | 2023-03-24 15:00 | 2023-03-24 16:00 | Maintenance | Prevent Maintenance | Maintenance |  | BME |  | Done |  | FALSE
59 | Bio-Immuno | Alinity C | AC-3 | 2023-03-23 11:40 |  | 2023-03-23 11:45 | 2023-03-23 16:30 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Vũ | BME | Done | Scheduled | FALSE
60 | Bio-Immuno | Alinity I | AI-3 | 2023-03-23 14:00 |  | 2023-03-23 11:45 | 2023-03-23 14:30 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Vũ | BME | Done | Scheduled | FALSE
61 | Bio-Immuno | Alinity I | AI-4 | 2023-03-23 14:00 |  | 2023-03-23 11:45 | 2023-03-23 14:30 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Vũ | BME | Done | Scheduled | FALSE
117 | ELISA | IMMUNOMAT | ELISA-04 | 2023-03-21 13:55 |  | 2023-03-21 14:00 | 2023-03-21 16:00 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Vinh | BME | Done | Scheduled | FALSE
118 | ELISA | IMMUNOMAT | ELISA-05 | 2023-03-21 13:55 |  | 2023-03-21 14:00 | 2023-03-21 16:00 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Vinh | BME | Done | Scheduled | FALSE
52 | GLP | SAL User Interface | LAS-UI3 | 2023-03-20 17:55 |  | 2023-03-20 18:00 | 2023-03-20 19:00 | Breakdown | The Barcode reader fail | Calibrate Barcode reader | Offline | Mr. Hoàng | BME | Done |  | FALSE
74 | GLP | SAL User Interface | LAS-UI1 | 2023-03-20 14:55 |  | 2023-03-20 15:00 | 2023-03-20 15:30 | Maintenance | Prevent Maintenance | Mantenance | Offline | Mr. Hoàng | BME | Done | Scheduled | FALSE
75 | GLP | SAL User Interface | LAS-UI2 | 2023-03-20 14:55 |  | 2023-03-20 15:00 | 2023-03-20 15:30 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Hoàng | BME | Done | Scheduled | FALSE
80 | GLP | Loader Module | LM-01 | 2023-03-20 11:55 |  | 2023-03-20 12:00 | 2023-03-20 13:30 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Hoàng | BME | Done | Scheduled | FALSE
83 | GLP | Buffer Module | BM-02 | 2023-03-20 11:55 |  | 2023-03-20 12:00 | 2023-03-20 13:00 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Hoàng | BME | Done | Scheduled | FALSE
85 | GLP | Input/ Output Module | IOM-02 | 2023-03-20 11:55 |  | 2023-03-20 12:00 | 2023-03-20 12:30 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Hoàng | BME | Done | Scheduled | FALSE
84 | GLP | Input/ Output Module | IOM-03 | 2023-03-20 11:55 |  | 2023-03-20 12:00 | 2023-03-20 12:30 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Hoàng | BME | Done | Scheduled | FALSE
174 | GLP | Rack Builder Module | RBM-C8000 | 2023-03-20 11:55 |  | 2023-03-20 12:00 | 2023-03-20 12:30 | Maintenance | Prevent Maintenance | Maintenane | Offline | Mr. Hoàng | BME | Done | Scheduled | FALSE
79 | GLP | Remover Cap Module | REM-01 | 2023-03-20 11:55 |  | 2023-03-20 12:00 | 2023-03-20 12:30 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Hoàng | BME | Done | Scheduled | FALSE
82 | GLP | Storage on SM | STORE-01 | 2023-03-20 11:55 |  | 2023-03-20 12:00 | 2023-03-20 12:30 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Hoàng | BME | Done | Scheduled | FALSE
81 | GLP | Aliquoter Module | AM-01 | 2023-03-20 10:55 |  | 2023-03-20 11:00 | 2023-03-20 11:30 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Hoàng | BME | Done | Scheduled | FALSE
76 | GLP | SAL User Interface | LAS-UI3 | 2023-03-20 10:55 |  | 2023-03-20 11:00 | 2023-03-20 11:30 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Hoàng | BME | Done | Scheduled | FALSE
51 | GLP | Cars and Track | CTR | 2023-03-18 10:55 |  | 2023-03-18 11:00 | 2023-03-18 12:00 | Breakdown | Weekly maintenance | Maintenacne |  | BME |  | Done |  | FALSE
53 | AO | Tempus | Tempus600 | 2023-03-17 13:55 |  | 2023-03-17 14:00 | 2023-03-17 14:30 | Maintenance | Prevent Maintenance | Clean Break part and change Rubber |  | BME |  | Done |  | FALSE
50 | Bio-Immuno | Immulite | MD-08 | 2023-03-16 13:55 |  | 2023-03-16 14:00 | 2023-03-16 17:00 | Breakdown | Reagent barcode scanning fail | Check and align Reagent barcode scanner | Offline | Mr. Thuận An | BME | Done |  | FALSE
126 | GLP | Bulk Loader Module | BLM | 2023-03-16 07:55 |  | 2023-03-16 08:00 | 2023-03-16 09:00 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Hoàng | BME | Done | Scheduled | FALSE
2305171622 | Bio-Immuno | Immulite | MD-08 | 2023-03-15 12:00 |  | 2023-03-15 13:26 | 2023-03-15 18:30 | Breakdown | Tube lifter failed | Change tube lifter and adjustment | Offline | Mr. Thụân An | BME | Done |  | FALSE |  |  | Form_Images/2305171622MD-08.Image.093112.jpg
46 | Bio-Immuno | Alinity C | AC-3 | 2023-03-15 16:00 |  | 2023-03-15 16:20 | 2023-03-15 16:40 | Breakdown | Needle's sample bended | Replace Needle's sample | Offline | Mr. Trúc | BME | Done |  | FALSE
128 | GLP | Tempus Module | IOM_TIO-02 | 2023-03-15 09:55 |  | 2023-03-15 10:00 | 2023-03-15 10:30 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Hoàng | BME | Done | Scheduled | FALSE
49 | Bio-Immuno | Alinity I | AI-1 | 2023-03-13 16:30 |  | 2023-03-13 17:00 | 2023-03-13 18:00 | Breakdown | Pretreatment path Stuck | Alignt pretreatment path, test load/unload RV | Offline | Mr. Vũ | BME | Done |  | TRUE | Vitamin B12: 2; Anti-HBc M: 2 | 4
47 | Bio-Immuno | Alinity C | AC-2 | 2023-03-13 14:00 |  | 2023-03-13 14:30 | 2023-03-13 14:50 | Breakdown | Sample's pipettor stuck | Change sample needle | Offline | Mr. Vũ | BME | Done |  | FALSE
2305180040 | Hematology - Coagulation | Alinity H-series | Hq-1 | 2023-03-13 08:40 |  | 2023-03-13 09:00 | 2023-03-13 09:50 | Breakdown | Aspiration probe out struction detected; tube bottom detected too soon; WBC temperature out of range | Change piece valve, reconnect heater cable | Offline | Mr. Thịnh | BME | Done |  | FALSE |  |  | Form_Images/2305180040Hq-1.Image.181754.png
39 | GLP | Cars and Track | CTR | 2023-03-11 13:55 |  | 2023-03-11 14:00 | 2023-03-11 15:00 | Maintenance | Prevent Maintenance | Clean all track element, check all database module after upgrade tsm 4.0 (tsb 330-009) | Offline | Mr. Hoàng | BME | Done | Scheduled | FALSE
108 | GLP | Stramax Connect Module | STR-01 | 2023-03-11 11:55 |  | 2023-03-11 12:00 | 2023-03-11 13:30 | Breakdown | AP can't connect | Change AP board | Offline | Mr. Tới | BME | Done |  | FALSE
40 | GLP | Cars and Track | CTR | 2023-03-11 10:55 |  | 2023-03-11 11:00 | 2023-03-11 12:00 | Maintenance | Prevent Maintenance | Maintenance |  | BME |  | Done |  | FALSE
48 | Bio-Immuno | Alinity C | AC-2 | 2023-03-11 11:00 |  | 2023-03-11 11:20 | 2023-03-11 11:45 | Breakdown | Sample needle bend | align sample needle | Offline | Mr. Vũ | BME | Done |  | TRUE | SCC: 1; Free PSA: 1; Anti-HBs: 1; Anti-HCVII: 1 | 4
109 | GLP | Tempus Module | IOM_TIO-02 | 2023-03-11 09:55 |  | 2023-03-11 10:00 | 2023-03-11 10:20 | Breakdown | Gripper finger stuck with sample tube | Take out the tube, clean and realign the gripper | Offline | BME |  | Done |  | FALSE
2305172259 | Hematology - Coagulation | Automated Blood typing machine | ABO-01 | 2023-03-10 10:00 |  | 2023-03-10 10:03 | 2023-03-10 11:30 | Breakdown | Create new test (ABO/Rh) | Create and active test | Offline | Mr. Hải | Mr. Sơn | Done |  | FALSE |  |  | Form_Images/2305172259ABO-01.Image.160335.png
66 | Microbiology | Urine Cell Analyzer UF-4000 | UF-01 | 2023-03-09 09:53 |  | 2023-03-09 13:52 | 2023-03-09 15:15 | Maintenance | Prevent Maintenance | maintenance | Offline | Mr. Đại | Ms. Hân | Done | Scheduled | FALSE
2304271516 | Hematology - Coagulation | Minicap Flex Piercing | DD-03 | 2023-03-08 09:06 |  | 2023-03-08 10:03 | 2023-03-08 16:22 | Breakdown | Pressure error | Check valve 8 ports, liquid valve, O-Ring --> valve fail | Offline | Mr. Huy | BME | Done | Cleaned probe; Waiting for change probe | FALSE
54 | Hematology - Coagulation | STA R MAX 3 | DM-03 | 2023-03-08 13:47 |  | 2023-03-08 13:48 | 2023-03-08 15:53 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Việt Anh | Mr. Trung | Done | Scheduled | FALSE
44 | GLP | Buffer Module | BM-01 | 2023-03-08 13:55 |  | 2023-03-08 14:00 | 2023-03-08 15:00 | Maintenance | Prevent Maintenance | clean and check robotic | Offline | Mr. Hoàng | BME | Done | Scheduled | FALSE
43 | GLP | Decaper Module Double | DM-01 | 2023-03-08 13:55 |  | 2023-03-08 14:00 | 2023-03-08 14:30 | Maintenance | Prevent Maintenance | Clean robotic, track element, check decapper finger and cable | Offline | Mr. Hoàng | BME | Done | Scheduled | FALSE
45 | GLP | Tube Assecssment Module | TAM | 2023-03-08 13:35 |  | 2023-03-08 13:40 | 2023-03-08 14:20 | Maintenance | Prevent maintenance | clean, check robotic, check TAC, re-align robotic and tac | Offline | Mr. Hoàng | BME | Done | Scheduled | FALSE
42 | GLP | Centrifuge Module | CM-03 | 2023-03-07 14:55 |  | 2023-03-07 15:00 | 2023-03-07 16:00 | Maintenance | Prevent Maintenance | clean, check and align robotic | Offline | Mr. Hoàng | BME | Done | Scheduled | FALSE
131 | GLP | Tube Assecssment Center | TAC1 | 2023-03-07 12:55 |  | 2023-03-07 13:00 | 2023-03-07 14:30 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Hoàng | BME | Done | Scheduled | FALSE
41 | GLP | Centrifuge Module | CM-01 | 2023-03-07 12:55 |  | 2023-03-07 13:00 | 2023-03-07 14:00 | Maintenance | Prevent Maintenance | Clean, check and align robotic | Offline | Mr. Hoàng | BME | Done | Scheduled | FALSE
72 | GLP | Centrifuge Module | CM-02 | 2023-03-07 12:55 |  | 2023-03-07 13:00 | 2023-03-07 14:00 | Maintenance | Prevent Maintenance | clean, check and align robotic | Offline | Mr. Hoàng | BME | Done | Scheduled | FALSE
37 | GLP | Cars and Track | CTR | 2023-03-05 11:25 |  | 2023-03-05 11:30 | 2023-03-05 12:30 | Breakdown | CrossSwitch damaged | Replace CrossSwitch | Offline | Mr. Hoàng | BME | Done |  | FALSE
2304271539 | Hematology - Coagulation | Alinity H-series | Hq-1 | 2023-03-04 09:30 |  | 2023-03-04 10:00 | 2023-03-04 | Breakdown | Fluid entered buffer tank | Clean valve 02, valve 07, waste pipe | Offline | Mr. Tới | BME | Done |  | FALSE
38 | GLP | Cars and Track | CTR | 2023-03-03 13:55 |  | 2023-03-03 14:00 | 2023-03-03 15:00 | Maintenance | Prevent Maintenance | Maintenance |  | BME |  | Done |  | FALSE
36 | GLP | Cars and Track | CTR | 2023-03-03 12:25 |  | 2023-03-03 12:30 | 2023-03-03 13:30 | Breakdown | CrossSwitch TAC failed | replace CrossSwitch | Offline | Mr. Hoàng | BME | Done |  | FALSE
127 | GLP | Tempus Module | IOM_TIO-01 | 2023-03-02 21:55 |  | 2023-03-02 22:00 | 2023-03-02 22:30 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Hoàng | BME | Done | Scheduled | FALSE
2305180231 | AO | Tempus | Tempus600 | 2023-03-01 15:31 |  | 2023-03-01 16:00 | 2023-03-01 19:00 | Breakdown | Dust on Tempus system | Clean pipe, receiving module, break system | Offline | Mr. Duy | BME | Follow up |  | FALSE |  |  | Form_Images/2305180231Tempus600.Image.193556.png
35 | GLP | Tempus Module | IOM_TIO-02 | 2023-03-01 09:35 |  | 2023-03-01 09:40 | 2023-03-01 11:30 | Breakdown | Module controller disconnected | Replace Module controller board | Offline | Mr. Hoàng | BME | Done | 178555657 5210006989 G-41013-01 MODULE CONTROLLER BOA A44 1 | FALSE
26 | Bio-Immuno | Cobas e801 Analytical unit | MD-09 | 2023-02-28 10:00 |  | 2023-02-28 10:20 | 2023-02-28 11:00 | Breakdown | Reagent loading error, lever caution | Check reagent pack, take log file for study | Offline | Mr. Phúc | Ms. Ngọc | Done | Waiting for answer after 2 weeks | FALSE
27 | GLP | Tempus Module | IOM_TIO-01 | 2023-02-27 23:55 |  | 2023-02-28 | 2023-02-28 00:15 | Breakdown | Tube out of line | Clean and research Drawer | Offline | Mr. Hoàng | BME | Done |  | FALSE
28 | GLP | RSM | RSM-2 | 2023-02-27 07:30 |  | 2023-02-27 07:45 | 2023-02-27 10:00 | Breakdown | 5684 RSMtransport pick error at sample positioner | Clean and lube axis, calibrate | Offline | Mr. Vũ | BME | Done |  | FALSE
56 | Hematology - Coagulation | Tosoh HLC-723 G11 LA analyzer | HbA1c-02 | 2023-02-24 15:00 |  | 2023-02-24 15:30 | 2023-02-24 16:00 | Breakdown | Prevent Maintenance | Maintenance | Offline | Mr. Trường | BME | Done | Scheduled | FALSE
55 | Hematology - Coagulation | Tosoh HLC-723 G11 analyzer | HbA1c-01 | 2023-02-24 14:55 |  | 2023-02-24 15:00 | 2023-02-24 15:30 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Trường | BME | Done | Scheduled | FALSE |  |  |  |  | Form_Images/2305221535QS7P-01.Image.084621.jpg
371 | Microbiology | Previ color gram | GR-01 | 2023-02-24 13:55 |  | 2023-02-24 14:00 | 2023-02-24 15:00 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Trường | BME | Done | Scheduled | FALSE
33 | GLP | Input/ Output Module | IOM-03 | 2023-02-24 08:55 |  | 2023-02-24 09:00 | 2023-02-24 09:20 | Breakdown | Module controller disconnected | Re-calibration rack port | Offline | Mr. Hoàng | BME | Done |  | FALSE
25 | GLP | Cars and Track | CTR | 2023-02-23 11:55 |  | 2023-02-23 12:00 | 2023-02-23 13:00 | Maintenance | Prevent Maintenance | Maintenance |  | BME |  | Done |  | FALSE
720 | GLP | Tosho Connect Module | TOS-01 | 2023-02-23 11:00 |  | 2023-02-23 11:05 | 2023-02-23 11:30 | Breakdown | No Host communication | Check and power up the Moxa connect | Offline | Mr. Trường | BME | Done | Power fail of Moxa connector | FALSE
34 | GLP | SAL User Interface | LAS-UI3 | 2023-02-23 09:55 |  | 2023-02-23 10:00 | 2023-02-23 10:20 | Breakdown | Sample Access Line (SAL); Tube Carrier; No Movement | Clean and reset AP, check config AP | Offline | Mr. Hoàng | BME | Done |  | FALSE
273 | Molecular Biology | Pipette 100-1000ul | FAM002303_2 | 2023-02-10 | 2023-02-10 | 2023-02-10 | 2023-02-22 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done | QH47409 | FALSE
987 | Molecular Biology | Pipette 1-10ul | FAM002305_2 | 2023-02-16 | 2023-02-16 | 2023-02-16 | 2023-02-22 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done | SH58787 | FALSE
502 | Molecular Biology | Pipette 1-10ul | FAM002305_3 | 2023-02-16 | 2023-02-16 | 2023-02-16 | 2023-02-22 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done | QH44336 | FALSE
460 | Molecular Biology | Pipette 20-200ul | FAM002306_1 | 2023-02-16 | 2023-02-16 | 2023-02-16 | 2023-02-22 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done | QH44335 | FALSE
594 | Molecular Biology | Pipette 20-200ul | FAM002306_2 | 2023-02-16 | 2023-02-16 | 2023-02-16 | 2023-02-22 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done | QH44334 | FALSE
651 | Molecular Biology | Pipette 20-200ul | FAM002306_3 | 2023-02-16 | 2023-02-16 | 2023-02-16 | 2023-02-22 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done | SH58789 | FALSE
129 | GLP | Stramax Connect Module | STR-01 | 2023-02-21 13:55 |  | 2023-02-21 14:00 | 2023-02-21 15:30 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Hoàng | BME | Done | Scheduled | FALSE
22 | AO | Tempus | Tempus600 | 2023-02-21 10:55 |  | 2023-02-21 11:00 | 2023-02-21 11:20 | Breakdown | Tube in Carousel, Cannot detect tube on TIM | Check Carousel, Check the Receiver Clean and bring online back | Online | Mr. Hoàng | BME | Done | Still have dust | FALSE
21 | Bio-Immuno | Cobas e801 Analytical unit | MD-09 | 2023-02-20 14:25 |  | 2023-02-20 14:30 | 2023-02-20 18:30 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Ngọc | BME | Done | Have to delay tests from 2:30 PM to 6 PM | FALSE
30 | Bio-Immuno | Alinity I | AI-1 | 2023-02-20 10:00 |  | 2023-02-20 10:35 | 2023-02-20 12:00 | Breakdown | 3424 (Sample) pipettor aspiration error | Clean sample probe, replace PM sensor, recalibration sample pipettor | Offline | Mr. Vũ | BME | Done |  | FALSE
2305180124 | Hematology - Coagulation | Alinity H-series | Hq-2 | 2023-02-20 08:30 |  | 2023-02-20 09:00 | 2023-02-20 11:30 | Breakdown | QC out of range | Change valve 02, clean sample probe wash area | Offline | Mr. Tới | Mr. Lâm | Done |  | FALSE |  |  | Form_Images/2305180124Hq-2.Image.182649.png
837 | Bio-Immuno | Pippette | PP-09 | 2023-02-16 | 2023-02-16 | 2023-02-16 | 2023-02-20 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done | Calibration outsite | FALSE
b6236648 | Bio-Immuno | RO system Roche | RO-MB | 2023-02-18 08:40 | 2023-02-18 08:41 | 2023-02-18 08:45 | 2023-02-18 09:45 | Maintenance | Prevent Maintenance | Change filter; Check power, water supply and transport | Offline | Mr. Tín | BME | Done |  | FALSE |  |  | report_Images/b6236648.Image.033849.png
4531 | Bio-Immuno | RO system Alinity | RO-MB | 2023-02-18 08:40 |  | 2023-02-18 08:45 | 2023-02-18 09:40 | Breakdown | Prevent Mainternance | Change Filter; Check power, water supply and transport | Offline | Mr. Tín | BME | Done |  | FALSE
29 | Bio-Immuno | Alinity C | AC-2 | 2023-02-16 17:25 |  | 2023-02-16 17:30 | 2023-02-16 22:30 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Vũ | BME | Done | Scheduled | FALSE
31 | Bio-Immuno | Alinity I | AI-2 | 2023-02-16 17:25 |  | 2023-02-16 17:30 | 2023-02-16 22:30 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Vũ | BME | Done | Scheduled | FALSE
20 | AO | Tempus | Tempus600 | 2023-02-15 22:00 |  | 2023-02-15 18:00 | 2023-02-15 19:00 | Breakdown | Dust on system, tube uncapped | Clean pipes, change line 1 to plastic pipe | Offline | Mr. Duy, Mr. Hoàng | BME | Done | If running stable can proceed with line 2/ Do clean when have no samples | FALSE
19 | GLP | Cars and Track | CTR | 2023-02-15 10:55 |  | 2023-02-15 11:00 | 2023-02-15 12:00 | Maintenance | Prevent Maintenance | Maintenance |  | BME |  | Done |  | FALSE
896 | AO | Centrifuge | LT-005 | 2023-02-11 | 2023-02-11 | 2023-02-11 | 2023-02-15 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE
265 | AO | Centrifuge | LT-006 | 2023-02-11 | 2023-02-11 | 2023-02-11 | 2023-02-15 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE
236 | Molecular Biology | Cold Centrifuge | LT-02 | 2023-02-11 | 2023-02-11 | 2023-02-11 | 2023-02-15 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE
584 | Molecular Biology | Agilent Microplate Centrifuge | MC-01 | 2023-02-11 | 2023-02-11 | 2023-02-11 | 2023-02-15 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE
507 | Microbiology | Shaker | ML-01 | 2023-02-11 | 2023-02-11 | 2023-02-11 | 2023-02-15 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE
392 | Molecular Biology | Minispin | MS-01 | 2023-02-11 | 2023-02-11 | 2023-02-11 | 2023-02-15 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE
996 | Molecular Biology | Multi-tube Vortexer | MVT-01 | 2023-02-11 | 2023-02-11 | 2023-02-11 | 2023-02-15 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE
23 | Molecular Biology | Multi-tube Vortexer | MVT-02 | 2023-02-11 | 2023-02-11 | 2023-02-11 | 2023-02-15 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE
733 | Molecular Biology | Mini vortexer | VT-01 | 2023-02-11 | 2023-02-11 | 2023-02-11 | 2023-02-15 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE
721 | Molecular Biology | Vortexter | VT-02 | 2023-02-11 | 2023-02-11 | 2023-02-11 | 2023-02-15 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE
141 | Bio-Immuno | Electronic hygrometer | AKĐT-1 | 2023-02-10 | 2023-02-10 | 2023-02-10 | 2023-02-14 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done | Calibration outsite | FALSE
157 | Warehouse | Electronic hygrometer | AKĐT-2 | 2023-02-10 | 2023-02-10 | 2023-02-10 | 2023-02-14 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done | Calibration outsite | FALSE
323 | Molecular Biology | Class II Biosafety Cabinet | BSC-02 | 2023-02-10 | 2023-02-10 | 2023-02-10 | 2023-02-14 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE
587 | Molecular Biology | Class II Biosafety Cabinet | BSC-03 | 2023-02-10 | 2023-02-10 | 2023-02-10 | 2023-02-14 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE
69 | Molecular Biology | Pipette 0.2-2ul | FAM002302_1 | 2023-02-10 | 2023-02-10 | 2023-02-10 | 2023-02-14 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done | SH78708 | FALSE
134 | Molecular Biology | Pipette 0.2-2ul | FAM002302_2 | 2023-02-10 | 2023-02-10 | 2023-02-10 | 2023-02-14 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done | SH56613 | FALSE
153 | Molecular Biology | Pipette 0.2-2ul | FAM002302_3 | 2023-02-10 | 2023-02-10 | 2023-02-10 | 2023-02-14 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done | QH47406 | FALSE
303 | Molecular Biology | Pipette 100-1000ul | FAM002303_1 | 2023-02-10 | 2023-02-10 | 2023-02-10 | 2023-02-14 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done | SH78709 | FALSE
647 | Molecular Biology | Pipette 1-10ul | FAM002305_1 | 2023-02-10 | 2023-02-10 | 2023-02-10 | 2023-02-14 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done | SH58790 | FALSE
251 | Bio-Immuno | Thermometer | NK-02 | 2023-02-10 | 2023-02-10 | 2023-02-10 | 2023-02-14 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done | Calibration outsite | FALSE
628 | Bio-Immuno | Thermometer | NK-11 | 2023-02-10 | 2023-02-10 | 2023-02-10 | 2023-02-14 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done | Calibration outsite | FALSE
920 | Warehouse | Electronic thermometer | NKĐT-1 | 2023-02-10 | 2023-02-10 | 2023-02-10 | 2023-02-14 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done | Calibration outsite | FALSE
705 | Bio-Immuno | Electronic thermometer | NKĐT-3 | 2023-02-10 | 2023-02-10 | 2023-02-10 | 2023-02-14 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done | Calibration outsite | FALSE
77 | Bio-Immuno | Electronic thermometer | NKĐT-4 | 2023-02-10 | 2023-02-10 | 2023-02-10 | 2023-02-14 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done | Calibration outsite | FALSE
912 | Bio-Immuno | Electronic thermometer | NKĐT-5 | 2023-02-10 | 2023-02-10 | 2023-02-10 | 2023-02-14 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done | Calibration outsite | FALSE
222 | Bio-Immuno | Electronic thermometer | NKĐT-6 | 2023-02-10 | 2023-02-10 | 2023-02-10 | 2023-02-14 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done | Calibration outsite | FALSE
910 | ELISA | Electronic thermometer | NKĐT-7 | 2023-02-10 | 2023-02-10 | 2023-02-10 | 2023-02-14 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done | Calibration outsite | FALSE
444 | Hematology - Coagulation | Electronic thermometer | NKĐT-8 | 2023-02-10 | 2023-02-10 | 2023-02-10 | 2023-02-14 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done | Calibration outsite | FALSE
422 | AO | Pipette | PP-06 | 2023-02-10 | 2023-02-10 | 2023-02-10 | 2023-02-14 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done | Calibration outsite | FALSE
765 | Hematology - Coagulation | Finnpipette | PP-10 | 2023-02-10 | 2023-02-10 | 2023-02-10 | 2023-02-14 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done | Calibration outsite | FALSE
302 | ELISA | Pipet | PP-13 | 2023-02-10 | 2023-02-10 | 2023-02-10 | 2023-02-14 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done | Calibration outsite | FALSE
388 | Hematology - Coagulation | Pipette Eppendorf | PP-14 | 2023-02-10 | 2023-02-10 | 2023-02-10 | 2023-02-14 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done | Calibration outsite | FALSE
238 | ELISA | Pippet 100 microlite (fixed) | PP-17 | 2023-02-10 | 2023-02-10 | 2023-02-10 | 2023-02-14 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done | Calibration outsite | FALSE
825 | Bio-Immuno | Pippette | PP-23 | 2023-02-10 | 2023-02-10 | 2023-02-10 | 2023-02-14 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done | Calibration outsite | FALSE
738 | Bio-Immuno | Pippette | PP-24 | 2023-02-10 | 2023-02-10 | 2023-02-10 | 2023-02-14 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done | Calibration outsite | FALSE
165 | Molecular Biology | Thermo-hygrometer | TH-01 | 2023-02-10 | 2023-02-10 | 2023-02-10 | 2023-02-14 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done | Calibration outsite | FALSE
815 | Molecular Biology | Thermo-hygrometer | TH-02 | 2023-02-10 | 2023-02-10 | 2023-02-10 | 2023-02-14 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done | Calibration outsite | FALSE
743 | Molecular Biology | Thermo-hygrometer | TH-03 | 2023-02-10 | 2023-02-10 | 2023-02-10 | 2023-02-14 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done | Calibration outsite | FALSE
326 | Molecular Biology | Thermo-hygrometer | TH-04 | 2023-02-10 | 2023-02-10 | 2023-02-10 | 2023-02-14 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done | Calibration outsite | FALSE
734 | Molecular Biology | Tempmate-M1 temperature data logger | TP-01 | 2023-02-10 | 2023-02-10 | 2023-02-10 | 2023-02-14 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done | Calibration outsite | FALSE
827 | Molecular Biology | Tempmate-M1 temperature data logger | TP-02 | 2023-02-10 | 2023-02-10 | 2023-02-10 | 2023-02-14 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done | Calibration outsite | FALSE
260 | Molecular Biology | Tempmate-M1 temperature data logger | TP-03 | 2023-02-10 | 2023-02-10 | 2023-02-10 | 2023-02-14 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done | Calibration outsite | FALSE
394 | Molecular Biology | Temperature USB Datalogger | TP-04 | 2023-02-10 | 2023-02-10 | 2023-02-10 | 2023-02-14 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done | Calibration outsite | FALSE
617 | Molecular Biology | Temperature USB Datalogger | TP-05 | 2023-02-10 | 2023-02-10 | 2023-02-10 | 2023-02-14 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done | Calibration outsite | FALSE
32 | Bio-Immuno | Alinity I | AI-2 | 2023-02-13 08:00 |  | 2023-02-13 08:15 | 2023-02-13 10:00 | Breakdown | Vacuum initialization failed | Replace drive board of vacuum pumb | Offline | Mr. Vũ | BME | Done |  | FALSE
581 | Molecular Biology | Refrigerator 2-8oC | CF-01 | 2023-02-09 | 2023-02-09 | 2023-02-09 | 2023-02-13 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE
379 | Molecular Biology | Refrigerator 2-8oC | CF-02 | 2023-02-09 | 2023-02-09 | 2023-02-09 | 2023-02-13 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE
294 | Molecular Biology | Refrigerator 2-8oC | CF-03 | 2023-02-09 | 2023-02-09 | 2023-02-09 | 2023-02-13 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE
57 | Molecular Biology | Refrigerator -20oC | DF-01 | 2023-02-09 | 2023-02-09 | 2023-02-09 | 2023-02-13 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE
247 | Molecular Biology | Refrigerator -70oC | DF-02 | 2023-02-09 | 2023-02-09 | 2023-02-09 | 2023-02-13 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE
101 | Warehouse | Cold warehouse | KHO-01 | 2023-02-09 | 2023-02-09 | 2023-02-09 | 2023-02-13 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE
792 | ELISA | Dry heat block machine | MU-01 | 2023-02-09 | 2023-02-09 | 2023-02-09 | 2023-02-13 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE
356 | ELISA | Dry heat block machine | MU-02 | 2023-02-09 | 2023-02-09 | 2023-02-09 | 2023-02-13 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE
462 | Molecular Biology | Autoclaves | NH-04 | 2023-02-09 | 2023-02-09 | 2023-02-09 | 2023-02-13 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE
546 | Molecular Biology | Water heating bath | NU-01 | 2023-02-09 | 2023-02-09 | 2023-02-09 | 2023-02-13 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE
179 | Bio-Immuno | Refrigerator | TL-005 | 2023-02-09 | 2023-02-09 | 2023-02-09 | 2023-02-13 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE
760 | AO | Refrigerator | TL-009 | 2023-02-09 | 2023-02-09 | 2023-02-09 | 2023-02-13 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE
158 | Bio-Immuno | Refrigerator -20oC | TL-06 | 2023-02-09 | 2023-02-09 | 2023-02-09 | 2023-02-13 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE
542 | Molecular Biology | Refrigerator -20oC | TL-07 | 2023-02-09 | 2023-02-09 | 2023-02-09 | 2023-02-13 | Calibration | Annual Calibration | Calibration | Offline | CTCR3 | BME | Done |  | FALSE
15 | AO | Tempus | Tempus600 | 2023-02-12 12:30 |  | 2023-02-12 13:00 | 2023-02-12 16:00 | Breakdown | Dust on system | Take out Line 1 pipe --> sent to the vendor for processing, clean system | Offline | Mr. Duy | BME | Done | Still have Line 2 for running | FALSE
2305180119 | Hematology - Coagulation | Alinity H-series | Hq-2 | 2023-02-11 08:20 |  | 2023-02-11 09:00 | 2023-02-11 12:30 | Breakdown | QC out of range | Check and clean dilute cup -> still out, recommend change QC | Offline | Mr. Thịnh | Mr. Lâm | Done |  | FALSE |  |  | Form_Images/2305180119Hq-2.Image.182253.png
16 | Microbiology | HeliFan Plus C13 | HP-01 | 2023-02-10 15:30 |  | 2023-02-10 16:00 | 2023-02-10 20:00 | Breakdown | Absorbed change -> Cannot zero gas, Absorder was expired | Change new absorber, clean aspirate pump, run system check, calibration | Offline | Mr. Vinh | BME | Done |  | TRUE | HP breath | 6
13 | GLP | Decaper Module Double | DM-01 | 2023-02-10 14:50 |  | 2023-02-10 15:00 | 2023-02-10 16:30 | Breakdown | Z axis collision failed | Decapper z axis down collision-->find out the flex cable is damaged FLEX CABLE 20X1040 -->replaced cable was fixed issue | Offline | Mr. Hoàng | BME | Done |  | FALSE
130 | GLP | Tosho Connect Module | TOS-01 | 2023-02-09 16:55 |  | 2023-02-09 17:00 | 2023-02-09 17:30 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Hoàng | BME | Done | Scheduled | FALSE
14 | GLP | Storage on SM | STORE-01 | 2023-02-08 15:00 |  | 2023-02-08 15:00 | 2023-02-08 16:00 | Breakdown | #21437 Storage init failed,near tower not as zero | Storage door can't reach to close position -->move it to close position and lubricant | Offline | Mr. Hoàng | BME | Done |  | FALSE
12 | GLP | Cars and Track | CTR | 2023-02-07 14:55 |  | 2023-02-07 15:00 | 2023-02-07 16:00 | Maintenance | Prevent Maintenance | Maintenance |  | BME |  | Done |  | FALSE
4 | AO | Tempus | Tempus600 | 2023-02-06 13:00 |  | 2023-02-06 13:30 | 2023-02-06 17:00 | Breakdown | Sensor receiving not match sending | Clean transfer pipe Keep Follow up | Offline | Mr. Hoàng | BME | Follow up |  | FALSE
932 | GLP | Storage on SM | Store-01 | 2023-02-03 10:25 |  | 2023-02-03 10:30 | 2023-02-03 11:00 | Breakdown | Griper stuck | Check and reset griper |  | BME |  | Done |  | FALSE
621 | GLP | Tempus Module | IOM_TIO-02 | 2023-02-02 10:25 |  | 2023-02-02 10:30 | 2023-02-02 11:30 | Breakdown | Tube out of line | Clean and recalib line | Offline | Mr. Hoàng | BME | Done |  | FALSE
430 | AO | Tempus | Tempus600 | 2023-02-02 09:55 |  | 2023-02-02 10:00 | 2023-02-02 10:40 | Breakdown | Tube stuck on carousel | Clean and re-Initialize carousel | Online | Mr. Hoàng | BME | Done |  | FALSE
3 | Bio-Immuno | Immulite | MD-08 | 2023-02-01 10:00 |  | 2023-02-01 11:00 | 2023-02-01 15:00 | Breakdown | Tube liter jam; System can not run autostart | General Check Check and align tube lifter Reinstall software, database backup | Offline | Mr. Thuận An | Ms. Huyền | Done |  | TRUE | FBC: 6 PAPP-A: 4 | 10
10 | GLP | Centrifuge Module | CM-01 | 2023-01-31 12:03 |  | 2023-01-31 12:08 | 2023-01-31 13:00 | Breakdown | Bucket not placed correctly | Realign CM robotic | Offline | Mr. Hoàng | BME | Done |  | FALSE
456 | GLP | Cars and Track | CTR | 2023-01-31 09:55 |  | 2023-01-31 10:00 | 2023-01-31 11:00 | Breakdown | Weekly Maintenance | Checking and Cleaning |  | BME |  | Done |  | FALSE
118 | AO | Tempus | Tempus600 | 2023-01-30 18:00 |  | 2023-01-30 18:45 | 2023-01-30 19:15 | Breakdown | Dust | Check pipe on Roof | Offline | Mr. Duy | BME | Done |  | FALSE
24 | GLP | Cars and Track | CTR | 2023-01-30 08:15 |  | 2023-01-30 08:20 | 2023-01-30 09:20 | Breakdown | Cross Switch do not response correctly | Change Cross Switch | Offline | Mr. Hoàng | BME | Done |  | FALSE
8 | Bio-Immuno | Alinity I | AI-4 | 2023-01-20 11:45 |  | 2023-01-20 11:50 | 2023-01-20 16:00 | Breakdown | Wash monitoring failed at Washzone 1 | Replace wz manifold, 2500ul pump, wz lift mechanism; Replace aspirate valve wzl, tighten fittings, clean wz waste manifold; Replace trigger distribution valve (valve not click); Replace trigger bypass valve (valve not click); Replace Temp optic board; Calibration and test wz1 -> pass | Offline | Mr. Vũ | BME | Done |  | TRUE | Toxo IgG: 5; SCC: 14; Rubbella IgM: 7; Cyffra 21-1: 2; CA 19-9 xR:6; CEA: 9; Anti HBe: 1 | 44
139 | AO | Tempus | Tempus600 | 2023-01-20 11:55 |  | 2023-01-20 12:00 | 2023-01-20 12:30 | Breakdown | Dust, stop machine | Stop and waiting to clean after Tet | Offline | Mr. Long | BME | Done |  | FALSE
921 | AO | Tempus | Tempus600 | 2023-01-20 07:55 |  | 2023-01-20 08:00 | 2023-01-20 08:20 | Breakdown | Stop working on receving modules | Clean sensor | Online | AO team | BME | Done |  | TRUE
378 | Hematology - Coagulation | Minicap Flex Piercing | DD-03 | 2023-01-18 09:00 |  | 2023-01-18 09:35 | 2023-01-18 16:09 | Breakdown | Tank broken and Probe stuck | Change tank to clean probe and wait to aprove for keep new tank | Offline | Mr. Huy | BME | Done |  | FALSE
9 | Bio-Immuno | Alinity C | AC-2 | 2023-01-18 08:40 |  | 2023-01-18 09:05 | 2023-01-18 12:15 | Breakdown | Reagent transport loading error | Take out reagent cartridges and realign reagent path | Offline | Mr. Vũ | BME | Done |  | FALSE
5 | Bio-Immuno | Alinity C | AC-1 | 2023-01-17 23:55 |  | 2023-01-18 | 2023-01-18 05:23 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Vũ | BME | Done | Scheduled | FALSE
6 | Bio-Immuno | Alinity I | AI-1 | 2023-01-17 23:55 |  | 2023-01-18 | 2023-01-18 03:01 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Vũ | BME | Done | Scheduled | FALSE
7 | Bio-Immuno | Alinity I | AI-4 | 2023-01-17 09:00 |  | 2023-01-17 09:15 | 2023-01-17 09:25 | Breakdown | Diluted Wash buffer transfer too slow | Replace the level sensor of diluted buffer | Offline | Mr. Vũ | BME | Done |  | FALSE
2 | Bio-Immuno | Immulite | MD-08 | 2023-01-16 13:00 |  | 2023-01-16 13:30 | 2023-01-16 18:40 | Breakdown | System can not run autostart | General check Reinstall the backup file, clean data Reinstall software v6.10 | Offline | Mr. Thuận An | BME | Done |  | FALSE
92 | GLP | Cars and Track | CTR | 2023-01-16 16:55 |  | 2023-01-16 17:00 | 2023-01-16 18:00 | Breakdown | Weekly Maintenance | Checking and Cleaning |  | BME |  | Done |  | FALSE
2305172354 | Hematology - Coagulation | Minicap Flex Piercing | DD-03 | 2023-01-14 08:42 |  | 2023-01-14 09:49 | 2023-01-14 13:35 | Breakdown | can't shutdown machine | Run backwash to check pressure -> failed -> Cathode tank broken -> glue | Offline | Mr. Huy | BME | Done | Waiting for change tank | FALSE |  |  | Form_Images/2305172354DD-03.Image.165748.png
606 | All | Sample Ladder | Lift | 2023-01-13 18:00 |  | 2023-01-13 18:45 | 2023-01-13 19:30 | Breakdown | Stuck, can't move | Clean and Check | Offline | Mr. Lâm | BME | Done |  | FALSE
1 | Bio-Immuno | Immulite | MD-08 | 2023-01-12 14:00 |  | 2023-01-12 15:00 | 2023-01-12 20:00 | Breakdown | Software error | Reinstall the backup file | Offline | Mr. Thuận An | BME | Done |  | TRUE | HCG: 6 | 6
349 | Molecular Biology | Multi-tube Vortexer | MVT-01 | 2023-01-12 15:55 |  | 2023-01-12 16:00 | 2023-01-12 18:00 | Breakdown | Loose shaft lock screw | Incresse size of screw thread and change screw snock |  | BME |  | Done |  | FALSE
364 | GLP | Loader Module | LM-01 | 2023-01-12 11:46 |  | 2023-01-12 11:51 | 2023-01-12 13:30 | Breakdown | Tube out of position when dispose | Clean and modify finger parameter |  | BME |  | Done |  | TRUE
597 | GLP | Cars and Track | CTR | 2023-01-11 15:55 |  | 2023-01-11 16:00 | 2023-01-11 17:00 | Breakdown | Weekly Maintenance | Checking and Cleaning |  | BME |  | Done |  | FALSE
965 | GLP | Loader Module | LM-01 | 2023-01-11 12:25 |  | 2023-01-11 12:30 | 2023-01-11 12:50 | Breakdown | Tube out of position when dispose | Clean area |  | BME |  | Done |  | FALSE | CBC | 75
277 | All | Attune | LIS | 2023-01-10 08:55 |  | 2023-01-10 09:00 | 2023-01-10 11:00 | Breakdown | CBC no order | Bring out and take test | Online | Mr. Kiên | BME | Done |  | TRUE
369 | AO | Tempus | Tempus600 | 2023-01-10 08:00 |  | 2023-01-10 08:30 | 2023-01-10 10:00 | Breakdown | Air Pressure failed | Check sensor | Offline | Mr. Quốc Anh | BME | Done |  | FALSE
174 | Bio-Immuno | RO system Abbott | RO-HP | 2023-01-09 13:55 |  | 2023-01-09 14:00 | 2023-01-09 15:30 | Breakdown | Maintenance | Change Presurre Increase Pumb | Offline | Mr. Tâm | BME | Done |  | FALSE
852 | AO | Tempus | Tempus600 | 2023-01-08 19:00 |  | 2023-01-08 21:00 | 2023-01-08 22:00 | Breakdown | Air Presurre Water Trap broke | Change Part | Offline | Mr. Quốc Anh | BME | Done | Tight fisrt and waiting for fixing | FALSE
17 | Cyto Lab | FUME HOOD | FH-01 | 2023-01-06 08:55 |  | 2023-01-06 09:00 | 2023-01-08 09:00 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Tài | Ms. Hà | Done | Scheduled | FALSE
18 | Cyto Lab | OVEN | OVEN-01 | 2023-01-06 08:55 |  | 2023-01-06 09:00 | 2023-01-06 09:30 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Tài | Ms. Hà | Done | Scheduled | FALSE
409 | AO | Tempus | Tempus600 | 2023-01-05 18:08 |  | 2023-01-05 18:13 | 2023-01-05 18:30 | Breakdown | Change Break's Rubber | Change Break's Rubber | Offline | Mr. Duy | BME | Done |  | FALSE
868 | GLP | Tempus Module | IOM_TIO-01 | 2023-01-04 21:55 |  | 2023-01-04 22:00 | 2023-01-04 22:15 | Breakdown | Collision of finger | Replace finger | Online | Mr. Hoàng | BME | Done |  | FALSE
570 | Hematology - Coagulation | Automated Blood typing machine | ABO-01 | 2023-01-04 13:55 |  | 2023-01-04 14:00 | 2023-01-04 16:00 | Breakdown | Maintenance | Prevent Maintenance | Offline | Mr. Điệu | BME | Done | Scheduled | FALSE
156 | GLP | Cars and Track | CTR | 2023-01-03 17:00 |  | 2023-01-03 17:00 | 2023-01-03 18:00 | Breakdown | Weekly Maintenance | Checking and Cleaning |  | BME |  | Done |  | FALSE
352 | Molecular Biology | Automated Laboratory Plate Sealer | ALPS | 2023-01-03 11:00 |  | 2023-01-03 11:00 | 2023-01-03 17:00 | Maintenance | Prevent Maintenance | Maintenance |  | BME |  | Done |  | FALSE
749 | Molecular Biology | #N/A | Ben | 2023-01-03 11:00 |  | 2023-01-03 11:00 | 2023-01-03 17:00 | Maintenance | Prevent Maintenance | Maintenance |  | BME |  | Done |  | FALSE
552 | Molecular Biology | KingFisher Flex | KFF-01 | 2023-01-03 11:00 |  | 2023-01-03 11:00 | 2023-01-03 17:00 | Maintenance | Prevent Maintenance | Maintenance |  | BME |  | Done |  | FALSE
522 | Molecular Biology | KingFisher Presto | KFP-01 | 2023-01-03 11:00 |  | 2023-01-03 11:00 | 2023-01-03 17:00 | Maintenance | Prevent Maintenance | Maintenance |  | BME |  | Done |  | FALSE
571 | Molecular Biology | KingFisher Presto | KFP-02 | 2023-01-03 11:00 |  | 2023-01-03 11:00 | 2023-01-03 17:00 | Maintenance | Prevent Maintenance | Maintenance |  | BME |  | Done |  | FALSE
909 | Molecular Biology | KingFisher Presto | KFP-03 | 2023-01-03 11:00 |  | 2023-01-03 11:00 | 2023-01-03 17:00 | Maintenance | Prevent Maintenance | Maintenance |  | BME |  | Done |  | FALSE
776 | Molecular Biology | KingFisher Presto | KFP-04 | 2023-01-03 11:00 |  | 2023-01-03 11:00 | 2023-01-03 17:00 | Maintenance | Prevent Maintenance | Maintenance |  | BME |  | Done |  | FALSE
571 | Molecular Biology | MultiDrop ComBI | MDC-01 | 2023-01-03 11:00 |  | 2023-01-03 11:00 | 2023-01-03 17:00 | Maintenance | Prevent Maintenance | Maintenance |  | BME |  | Done |  | FALSE
338 | Molecular Biology | MultiDrop ComBI | MDC-02 | 2023-01-03 11:00 |  | 2023-01-03 11:00 | 2023-01-03 17:00 | Maintenance | Prevent Maintenance | Maintenance |  | BME |  | Done |  | FALSE
518 | Molecular Biology | MultiDrop ComBI | MDC-03 | 2023-01-03 11:00 |  | 2023-01-03 11:00 | 2023-01-03 17:00 | Maintenance | Prevent Maintenance | Maintenance |  | BME |  | Done |  | FALSE
644 | Molecular Biology | Microplate Washer | MPW-01 | 2023-01-03 11:00 |  | 2023-01-03 11:00 | 2023-01-03 17:00 | Maintenance | Prevent Maintenance | Maintenance |  | BME |  | Done |  | FALSE
923 | Molecular Biology | Spinnaker Microplate Robot | Spinnaker | 2023-01-03 11:00 |  | 2023-01-03 11:00 | 2023-01-03 17:00 | Maintenance | Prevent Maintenance | Maintenance |  | BME |  | Done |  | FALSE

## Sheet1
Instrument Breakdown
Month | 7
ID | Department | Name | Instrument ID | Instrument Breakdown Date | Contact Time | Suport Time | Complete Time | Type | Detail | Action | Offline/ Online Support | Suport Persion | Contact Persion | Current Status | Note | Test affected during breakdown (yes/No) | Name of the test affected | No of test affected | Image | File
63ca6d1b | GLP | Tube Assecssment Module | TAM | 2023-07-28 11:57 | 2023-07-28 11:57 | 2023-07-28 11:59 | 2023-07-28 12:11 | Breakdown | Finger cut through tube cap | Take out, clean, re initialize | Offline | BME | BME | Done | 0 | FALSE | 0 | 0 | report_Images/63ca6d1b.Image.051649.jpg | 0
31 | GLP | Cars and Track | CTR | 2023-07-28 11:00 | 2023-07-28 11:00 | 2023-07-28 11:00 | 2023-07-28 12:00 | Maintenance | Prevent maintenance | Maintenance | Offline | BME | 0 | Done | 0 | FALSE | 0 | 0 | 0 | 0
9d704970 | GLP | SAL User Interface | LAS-UI1 | 2023-07-27 09:00 | 2023-07-27 09:00 | 2023-07-27 09:00 | 2023-07-27 10:00 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Hoàng | BME | Done | 0 | FALSE | 0 | 0 | report_Images/9d704970.Image.093305.png | report_Files_/9d704970.File.093305.pdf
cf93008a | GLP | SAL User Interface | LAS-UI2 | 2023-07-27 09:00 | 2023-07-27 09:00 | 2023-07-27 09:00 | 2023-07-27 10:00 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Hoàng | BME | Done | 0 | FALSE | 0 | 0 | report_Images/cf93008a.Image.093735.png | report_Files_/cf93008a.File.093735.pdf
d0ebf150 | GLP | SAL User Interface | LAS-UI3 | 2023-07-27 09:00 | 2023-07-27 09:00 | 2023-07-27 09:00 | 2023-07-27 10:00 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Hoàng | BME | Done | 0 | FALSE | 0 | 0 | report_Images/d0ebf150.Image.093414.png | report_Files_/d0ebf150.File.093414.pdf
e7492a43 | GLP | Rack Builder Module | RBM-C8000 | 2023-07-27 09:00 | 2023-07-27 09:00 | 2023-07-27 09:00 | 2023-07-27 10:00 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Hoàng | BME | Done | 0 | FALSE | 0 | 0 | report_Images/e7492a43.Image.093613.png | report_Files_/e7492a43.File.093613.pdf
076608ae | GLP | Storage on SM | STORE-01 | 2023-07-27 09:00 | 2023-07-27 09:00 | 2023-07-27 09:00 | 2023-07-27 10:00 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Hoàng | BME | Done | 0 | FALSE | 0 | 0 | report_Images/076608ae.Image.093101.png | report_Files_/076608ae.File.093101.pdf
0be204a6 | GLP | Tosho Connect Module | TOS-01 | 2023-07-27 09:00 | 2023-07-27 09:00 | 2023-07-27 09:00 | 2023-07-27 10:00 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Hoàng | BME | Done | 0 | FALSE | 0 | 0 | report_Images/0be204a6.Image.090323.png | report_Files_/0be204a6.File.090323.pdf
756ddbe6 | GLP | Bulk Loader Module | BLM | 2023-07-26 09:00 | 2023-07-26 09:00 | 2023-07-26 09:00 | 2023-07-26 10:00 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Hoàng | BME | Done | 0 | FALSE | 0 | 0 | report_Images/756ddbe6.Image.092224.png | report_Files_/756ddbe6.File.092224.pdf
b601195e | GLP | Decaper Module Double | DM-01 | 2023-07-26 09:00 | 2023-07-26 09:00 | 2023-07-26 09:00 | 2023-07-26 10:00 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Hoàng | BME | Done | 0 | FALSE | 0 | 0 | report_Images/b601195e.Image.092601.png | report_Files_/b601195e.File.092601.pdf
f7087bf7 | GLP | Tempus Module | IOM_TIO-01 | 2023-07-26 09:00 | 2023-07-26 09:00 | 2023-07-26 09:00 | 2023-07-26 10:00 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Hoàng | BME | Done | 0 | FALSE | 0 | 0 | report_Images/f7087bf7.Image.091733.png | report_Files_/f7087bf7.File.091733.pdf
133cba5e | GLP | Tempus Module | IOM_TIO-02 | 2023-07-26 09:00 | 2023-07-26 09:00 | 2023-07-26 09:00 | 2023-07-26 10:00 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Hoàng | BME | Done | 0 | FALSE | 0 | 0 | report_Images/133cba5e.Image.092103.png | report_Files_/133cba5e.File.092103.pdf
78d7cb02 | GLP | Loader Module | LM-01 | 2023-07-26 09:00 | 2023-07-26 09:00 | 2023-07-26 09:00 | 2023-07-26 10:00 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Hoàng | BME | Done | 0 | FALSE | 0 | 0 | report_Images/78d7cb02.Image.091922.png | report_Files_/78d7cb02.File.091922.pdf
9278bdc6 | GLP | Remover Cap Module | REM-01 | 2023-07-26 09:00 | 2023-07-26 09:00 | 2023-07-26 09:00 | 2023-07-26 10:00 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Hoàng | BME | Done | 0 | FALSE | 0 | 0 | report_Images/9278bdc6.Image.092943.png | report_Files_/9278bdc6.File.092943.pdf
b73351a2 | GLP | Recapper Module | RM-01 | 2023-07-26 09:00 | 2023-07-26 09:00 | 2023-07-26 09:00 | 2023-07-26 10:00 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Hoàng | BME | Done | 0 | FALSE | 0 | 0 | report_Images/b73351a2.Image.092813.png | report_Files_/b73351a2.File.092813.pdf
866fe204 | AO | Tempus | Tempus600 | 2023-07-25 10:51 | 2023-07-25 10:52 | 2023-07-25 10:53 | 2023-07-25 11:10 | Breakdown | Stuck sample in break | Take out and re-intialization | Offline | BME | BME | Follow up | 0 | FALSE | 0 | 0 | 0 | 0
6e098c95 | GLP | Aliquoter Module | AM-01 | 2023-07-25 09:00 | 2023-07-25 09:00 | 2023-07-25 09:00 | 2023-07-25 10:00 | Maintenance | Prevent Maintenance | 0 | Offline | Mr. Hoàng | BME | Done | 0 | FALSE | 0 | 0 | report_Images/6e098c95.Image.102002.png | report_Files_/6e098c95.File.102002.pdf
8e91e0f2 | GLP | Centrifuge Module | CM-01 | 2023-07-25 09:00 | 2023-07-25 09:00 | 2023-07-25 09:00 | 2023-07-25 10:00 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Hoàng | BME | Done | 0 | FALSE | 0 | 0 | report_Images/8e91e0f2.Image.092324.png | report_Files_/8e91e0f2.File.092324.pdf
88539b08 | GLP | Centrifuge Module | CM-02 | 2023-07-25 09:00 | 2023-07-25 09:00 | 2023-07-25 09:00 | 2023-07-25 10:00 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Hoàng | BME | Done | 0 | FALSE | 0 | 0 | report_Images/88539b08.Image.091014.png | 0
64b5d1a8 | GLP | Centrifuge Module | CM-03 | 2023-07-25 09:00 | 2023-07-25 09:00 | 2023-07-25 09:00 | 2023-07-25 10:00 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Hoàng | BME | Done | 0 | FALSE | 0 | 0 | report_Images/64b5d1a8.Image.090916.png | report_Files_/64b5d1a8.File.090916.pdf
097a077c | GLP | Tube Assecssment Module | TAM | 2023-07-25 09:00 | 2023-07-25 09:00 | 2023-07-25 09:00 | 2023-07-25 10:00 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Hoàng | BME | Done | 0 | FALSE | 0 | 0 | report_Images/097a077c.Image.091350.png | report_Files_/097a077c.File.091350.pdf
24ba6eb2 | GLP | Stramax Connect Module | STR-01 | 2023-07-24 13:52 | 2023-07-24 14:02 | 2023-07-24 14:32 | 2023-07-24 15:32 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Hoàng | BME | Done | 0 | FALSE | 0 | 0 | 0 | report_Files_/24ba6eb2.File.083320.pdf
b1f05e86 | GLP | Input/ Output Module | IOM-03 | 2023-07-24 13:41 | 2023-07-24 13:51 | 2023-07-24 14:21 | 2023-07-24 15:21 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Hoàng | BME | Done | 0 | FALSE | 0 | 0 | report_Images/b1f05e86.Image.082340.png | report_Files_/b1f05e86.File.083104.pdf
700858bc | GLP | Buffer Module | BM-01 | 2023-07-24 08:30 | 2023-07-24 08:30 | 2023-07-24 08:30 | 2023-07-24 10:30 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Hoàng | BME | Done | 0 | FALSE | 0 | 0 | report_Images/700858bc.Image.081807.png | 0
3d4e650c | GLP | Buffer Module | BM-02 | 2023-07-24 08:30 | 2023-07-24 08:30 | 2023-07-24 08:30 | 2023-07-24 10:00 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Hoàng | BME | Done | 0 | FALSE | 0 | 0 | report_Images/3d4e650c.Image.081555.png | report_Files_/3d4e650c.File.083736.pdf
05dd9c15 | ELISA | DAR800 | ELISA-02 | 2023-07-21 14:10 | 2023-07-21 14:10 | 2023-07-21 14:20 | 2023-07-21 16:00 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Bình | BME | Done | 0 | FALSE | 0 | 0 | 0 | 0
3fcc67b3 | Microbiology | HeliFan Plus C13 | HP-01 | 2023-07-21 14:10 | 2023-07-21 14:10 | 2023-07-21 14:20 | 2023-07-21 16:00 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Vinh | BME | Done | 0 | FALSE | 0 | 0 | 0 | 0
ccb66594 | ELISA | DAW50 | ELISA-01 | 2023-07-21 14:10 | 2023-07-21 14:10 | 2023-07-21 14:20 | 2023-07-21 15:52 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Bình | BME | Done | 0 | FALSE | 0 | 0 | 0 | 0
7bb4ace6 | ELISA | DAW50TS | ELISA-03 | 2023-07-21 14:10 | 2023-07-21 14:10 | 2023-07-21 14:20 | 2023-07-21 15:52 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Bình | BME | Done | 0 | FALSE | 0 | 0 | 0 | 0
78583fa6 | GLP | Input/ Output Module | IOM-02 | 2023-07-21 13:57 | 2023-07-21 14:07 | 2023-07-21 14:37 | 2023-07-21 15:37 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Hoàng | BME | Done | 0 | FALSE | 0 | 0 | 0 | report_Files_/78583fa6.File.083801.pdf
a4ec4c6d | GLP | Input/ Output Module | IOM-02 | 2023-07-21 13:55 | 2023-07-21 14:05 | 2023-07-21 14:35 | 2023-07-21 15:35 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Hoàng | BME | Done | 0 | FALSE | 0 | 0 | 0 | report_Files_/a4ec4c6d.File.083539.pdf
8e88b205 | GLP | Tube Assecssment Center | TAC2 | 2023-07-21 13:53 | 2023-07-21 14:03 | 2023-07-21 14:33 | 2023-07-21 15:33 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Hoàng | BME | Done | 0 | FALSE | 0 | 0 | 0 | report_Files_/8e88b205.File.083448.pdf
957b4455 | Hematology - Coagulation | STA R MAX 3 | DM-03 | 2023-07-20 15:50 | 2023-07-20 16:00 | 2023-07-20 16:05 | 2023-07-20 16:30 | Breakdown | Stucked cuvette | Clean cuvette pusher | Online | Mr. Huy | BME | Done | 0 | FALSE | 0 | 0 | 0 | 0
7a1da242 | Cyto Lab | FUME HOOD | FH-01 | 2023-07-18 16:56 | 2023-07-18 16:56 | 2023-07-18 16:56 | 2023-07-18 17:36 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Tài | BME | Done | 0 | FALSE | 0 | 0 | 0 | 0
eac7e26e | Cyto Lab | OVEN | OVEN-01 | 2023-07-18 16:56 | 2023-07-18 16:56 | 2023-07-18 16:56 | 2023-07-18 17:36 | Maintenance | Prevent Maintenance | Maintenance | Offline | Mr. Tài | BME | Done | 0 | FALSE | 0 | 0 | 0 | 0
c19e1b3c | Bio-Immuno | Alinity C | AC-2 | 2023-07-18 14:01 | 2023-07-18 14:02 | 2023-07-18 14:05 | 2023-07-18 14:20 | Breakdown | Optic error | Change lamp | Online | Mr. Vũ | BME | Done | 0 | FALSE | 0 | 0 | 0 | 0
34513 | GLP | Cars and Track | CTR | 2023-07-14 14:00 | 2023-07-14 14:00 | 2023-07-14 14:00 | 2023-07-14 15:00 | Maintenance | Prevent maintenance | Maintenance | Offline | BME | 0 | Done | 0 | FALSE | 0 | 0 | 0 | 0
88b44636 | AO | Tempus | Tempus600 | 2023-07-14 08:15 | 2023-07-14 08:18 | 2023-07-14 08:48 | 2023-07-14 09:48 | Breakdown | Tube not send on Carousel | Clean sensor, carousel | Offline | BME | BME | Done | 0 | FALSE | 0 | 0 | 0 | 0
78dd6b7a | All | Sample Ladder | Lift | 2023-07-13 13:46 | 2023-07-13 13:56 | 2023-07-13 14:26 | 2023-07-13 15:26 | Breakdown | counterweight rope break | Connect cable | Offline | Mr. Lâm | BME | Done | 0 | FALSE | 0 | 0 | 0 | 0
a30d16c6 | Bio-Immuno | Alinity I | AI-2 | 2023-07-13 11:43 | 2023-07-13 11:53 | 2023-07-13 12:00 | 2023-07-13 12:23 | Breakdown | Board communication error | Change fuse | Offline | Mr. Hoàng | BME | Done | 0 | FALSE | 0 | 0 | 0 | 0
86aa6b27 | AO | Tempus | Tempus600 | 2023-07-12 16:27 | 2023-07-13 10:37 | 2023-07-13 11:07 | 2023-07-13 12:07 | Breakdown | Sending sensor error | Clean sensor and run test Programme | Offline | Mr. Hoàng | BME | Follow up | 0 | FALSE | 0 | 0 | report_Images/86aa6b27.Image.050912.jpg | 0
c3a25f1f | Bio-Immuno | Alinity I | AI-1 | 2023-07-08 09:43 | 2023-07-08 09:45 | 2023-07-08 10:15 | 2023-07-08 10:43 | Breakdown | Pre-trigger level sensor error | Change level sensor | Offline | Mr. Tuấn Anh | Ms. Trân | Follow up | 0 | FALSE | 0 | 0 | 0 | 0
46 | GLP | Cars and Track | CTR | 2023-07-07 14:00 | 2023-07-07 14:00 | 2023-07-07 14:00 | 2023-07-07 15:00 | Maintenance | Prevent maintenance | Maintenance | Offline | BME | 0 | Done | 0 | FALSE | 0 | 0 | 0 | 0
13f9b0bf | GLP | Centrifuge Module | CM-01 | 2023-07-06 15:00 | 2023-07-06 15:01 | 2023-07-06 15:05 | 2023-07-06 15:15 | Breakdown | Finger cut through tube cap | Take out all tubes; clean and re init machine | Offline | BME | 0 | Done | 0 | FALSE | 0 | 0 | report_Images/13f9b0bf.Image.081909.jpg | 0
2fb942ae | Hematology - Coagulation | Tosoh HLC-723 G11 LA analyzer | HbA1c-02 | 2023-07-06 10:07 | 2023-07-06 10:17 | 2023-07-06 10:47 | 2023-07-06 11:47 | Breakdown | Valve have error | Check liquid line, valve and restart | Online | Mr. Trường | Mr. Lâm | Follow up | 0 | FALSE | 0 | 0 | 0 | 0
ef83697a | Hematology - Coagulation | Alinity H-series | Hq-1 | 2023-07-03 11:29 | 2023-07-03 11:39 | 2023-07-03 12:09 | 2023-07-03 13:09 | Breakdown | WBC not normal | Clean WBC tank | Offline | Mr. Tới | BME | Follow up | 0 | FALSE | 0 | 0 | 0 | 0

## Sheet2
Count of Detail | Column Labels
| 3-Jan | 4-Jan | 5-Jan | 6-Jan | 8-Jan | 9-Jan | 10-Jan | 11-Jan | 12-Jan | 13-Jan | 14-Jan | 16-Jan | 17-Jan | 18-Jan | 20-Jan | 30-Jan | 31-Jan | 1-Feb | 2-Feb | 3-Feb | 6-Feb | 7-Feb | 8-Feb | 9-Feb | 10-Feb | 11-Feb | 12-Feb | 13-Feb | 14-Feb | 15-Feb | 16-Feb | 18-Feb | 20-Feb | 21-Feb | 22-Feb | 23-Feb | 24-Feb | 27-Feb | 28-Feb | 1-Mar | 2-Mar | 3-Mar | 4-Mar | 5-Mar | 7-Mar | 8-Mar | 9-Mar | 10-Mar | 11-Mar | 13-Mar | 15-Mar | 16-Mar | 17-Mar | 18-Mar | 20-Mar | 21-Mar | 23-Mar | 24-Mar | 27-Mar | 28-Mar | 29-Mar | 30-Mar | 31-Mar | 3-Apr | 4-Apr | 5-Apr | 6-Apr | 7-Apr | 8-Apr | 9-Apr | 11-Apr | 12-Apr | 13-Apr | 19-Apr | 20-Apr | 21-Apr | 24-Apr | 25-Apr | 26-Apr | 27-Apr | 28-Apr | 3-May | 4-May | 5-May | 9-May | 10-May | 11-May | 12-May | 13-May | 16-May | 17-May | 18-May | 19-May | 22-May | 23-May | 24-May | 25-May | 26-May | 27-May | 29-May | 30-May | 2-Jun | 6-Jun | 7-Jun | 12-Jun | 15-Jun | 16-Jun | 19-Jun | 20-Jun | 21-Jun | 22-Jun | 23-Jun | 30-Jun | 3-Jul | 6-Jul | 7-Jul | 8-Jul | 13-Jul | 14-Jul | 18-Jul | 20-Jul | 21-Jul | 24-Jul | 25-Jul | 28-Jul | 1-Aug | Grand Total
Row Labels
ABI-04 | [67 empty] | 1 | [58 empty] | 1
ABO-01 |  | 1 | [45 empty] | 1 | [51 empty] | 2 | [26 empty] | 4
AC-1 | [13 empty] | 1 | [58 empty] | 1 | [21 empty] | 1 | [31 empty] | 3
AC-2 | [13 empty] | 1 | [16 empty] | 1 | [17 empty] | 1 | 1 | [22 empty] | 1 | [46 empty] | 1 | [6 empty] | 6
AC-3 | [50 empty] | 1 | [5 empty] | 1 | [15 empty] | 1 | 1 | [7 empty] | 1 | 1 |  |  |  | 1 | [39 empty] | 7
AI-1 | [13 empty] | 1 | [18 empty] | 1 | [16 empty] | 1 | [44 empty] | 1 | [21 empty] | 1 | [9 empty] | 5
AI-2 | [27 empty] | 1 |  |  | 1 | [58 empty] | 1 | [18 empty] | 1 | [8 empty] | 1 | [8 empty] | 5
AI-3 | [56 empty] | 1 | [17 empty] | 1 | [14 empty] | 1 | [36 empty] | 3
AI-4 | [12 empty] | 1 |  | 1 | [41 empty] | 1 | [10 empty] | 1 | [58 empty] | 4
AK-01 | [89 empty] | 1 | [36 empty] | 1
AK-03 | [89 empty] | 1 | [36 empty] | 1
AK-04 | [89 empty] | 1 | [36 empty] | 1
AK-05 | [89 empty] | 1 | [36 empty] | 1
AKĐT-1 | [28 empty] | 1 | [97 empty] | 1
AKĐT-2 | [28 empty] | 1 | [97 empty] | 1
ALPS | 1 | [110 empty] | 1 | [14 empty] | 2
AM-01 | [54 empty] | 1 | [38 empty] | 1 | [32 empty] | 2
ATSH-02 | [80 empty] | 1 | [45 empty] | 1
ATSH-03 | [80 empty] | 1 | [45 empty] | 1
ATSH-04 | [80 empty] | 1 | [45 empty] | 1
Ben | [111 empty] | 1 | [14 empty] | 1
Ben | 1 | [125 empty] | 1
BLM | [51 empty] | 1 | [74 empty] | 1
BM-01 | [45 empty] | 1 | [76 empty] | 1 |  |  |  | 2
BM-02 | [54 empty] | 1 | [67 empty] | 1 |  |  |  | 2
BSC-01 | [80 empty] | 1 | [45 empty] | 1
BSC-02 | [28 empty] | 1 | [97 empty] | 1
BSC-03 | [28 empty] | 1 | [97 empty] | 1
CAN-01 | [80 empty] | 2 | [45 empty] | 2
CF-01 | [27 empty] | 1 | [98 empty] | 1
CF-02 | [27 empty] | 1 | [98 empty] | 1
CF-03 | [27 empty] | 1 | [98 empty] | 1
CM-01 | [16 empty] | 1 | [27 empty] | 1 | [69 empty] | 1 | [11 empty] | 3
CM-02 | [44 empty] | 1 | [13 empty] | 1 | [67 empty] | 2
CM-03 | [44 empty] | 1 | [33 empty] | 1 | 1 | [13 empty] | 2 | [32 empty] | 5
Core 4 | [97 empty] | 1 | [28 empty] | 1
CTR | 1 | [6 empty] | 1 |  |  |  | 1 |  |  |  | 1 | 1 |  |  |  |  | 1 | [7 empty] | 1 | [5 empty] | 1 | [5 empty] | 2 |  | 1 |  |  |  |  | 2 |  |  |  |  | 1 |  |  |  | 1 | [5 empty] | 1 | 1 |  | 1 |  |  |  |  | 1 |  |  |  |  | 1 | [5 empty] | 1 |  |  | 1 |  | 1 |  |  |  |  | 1 |  |  |  |  | 1 |  |  |  | 1 |  |  | 1 |  | 1 |  |  |  |  | 1 | 1 |  |  | 1 |  |  | 1 | [5 empty] | 1 | 1 | 34
DD-03 | [10 empty] | 1 |  |  | 1 | [31 empty] | 1 | [29 empty] | 1 | 1 |  |  | 1 | [46 empty] | 6
DF-01 | [27 empty] | 1 | [98 empty] | 1
DF-02 | [27 empty] | 1 | [98 empty] | 1
DM-01 | [24 empty] | 1 | [20 empty] | 1 | [56 empty] | 1 | [23 empty] | 3
DM-03 | [45 empty] | 1 | [74 empty] | 1 | [5 empty] | 2
ELISA-01 | [61 empty] | 1 | [59 empty] | 1 |  |  |  |  | 2
ELISA-02 | [61 empty] | 1 | [23 empty] | 1 | [35 empty] | 1 |  |  |  |  | 3
ELISA-03 | [61 empty] | 1 | [59 empty] | 1 |  |  |  |  | 2
ELISA-04 | [55 empty] | 1 | [70 empty] | 1
ELISA-05 | [55 empty] | 1 | [70 empty] | 1
F2-PP10 | [89 empty] | 1 | [36 empty] | 1
F2-PP1000 | [89 empty] | 1 | [36 empty] | 1
F2-PP200 | [89 empty] | 1 | [36 empty] | 1
FAM002302_1 | [28 empty] | 1 | [97 empty] | 1
FAM002302_2 | [28 empty] | 1 | [97 empty] | 1
FAM002302_3 | [28 empty] | 1 | [97 empty] | 1
FAM002303_1 | [28 empty] | 1 | [97 empty] | 1
FAM002303_2 | [34 empty] | 1 | [91 empty] | 1
FAM002305_1 | [28 empty] | 1 | [97 empty] | 1
FAM002305_2 | [34 empty] | 1 | [91 empty] | 1
FAM002305_3 | [34 empty] | 1 | [91 empty] | 1
FAM002306_1 | [34 empty] | 1 | [91 empty] | 1
FAM002306_2 | [34 empty] | 1 | [91 empty] | 1
FAM002306_3 | [34 empty] | 1 | [91 empty] | 1
FH-01 |  |  |  |  | 1 | [61 empty] | 1 | [13 empty] | 1 | [38 empty] | 1 | [6 empty] | 4
GNV-01 | [77 empty] | 1 | [48 empty] | 1
GR-01 | [36 empty] | 1 | [68 empty] | 1 | [20 empty] | 2
HbA1c-01 | [36 empty] | 1 | [68 empty] | 1 | [20 empty] | 2
HbA1c-02 | [36 empty] | 1 | [36 empty] | 1 |  |  | 1 |  | 1 | [26 empty] | 1 | [8 empty] | 1 | [11 empty] | 6
HP-01 | [24 empty] | 1 | [36 empty] | 1 | [59 empty] | 1 |  |  |  |  | 3
Hq-1 | [42 empty] | 1 | [6 empty] | 1 | [11 empty] | 1 | [7 empty] | 1 | 1 |  | 1 | [23 empty] | 1 |  |  |  |  | 1 | [7 empty] | 1 |  | 1 |  | 1 | [12 empty] | 11
Hq-2 | [25 empty] | 1 | [6 empty] | 1 | [28 empty] | 2 | [43 empty] | 1 |  | 1 |  | 1 |  | 1 | [14 empty] | 8
IOM_TIO-01 |  | 1 | [36 empty] | 1 |  | 1 | [18 empty] | 1 | [7 empty] | 1 | [16 empty] | 1 | [5 empty] | 1 | [7 empty] | 1 | [5 empty] | 1 | [5 empty] | 1 | [15 empty] | 10
IOM_TIO-02 | [18 empty] | 1 | [20 empty] | 1 | [8 empty] | 1 |  | 1 | [8 empty] | 1 | [7 empty] | 1 | 1 | [15 empty] | 1 | [5 empty] | 1 |  | 1 | [11 empty] | 1 | [21 empty] | 11
IOM-02 | [54 empty] | 1 | [66 empty] | 2 |  |  |  |  | 3
IOM-03 | [36 empty] | 1 | [17 empty] | 1 | [30 empty] | 1 | [36 empty] | 1 |  |  |  | 4
KFF-01 | 1 | [60 empty] | 1 | 1 | [63 empty] | 3
KFP-01 | 1 | [110 empty] | 1 | [14 empty] | 2
KFP-02 | 1 | [110 empty] | 1 | [14 empty] | 2
KFP-03 | 1 | [110 empty] | 1 | [14 empty] | 2
KFP-04 | 1 | [110 empty] | 1 | [14 empty] | 2
KHO-01 | [27 empty] | 1 | [98 empty] | 1
KHV-01 | [62 empty] | 1 | [63 empty] | 1
KHV-02 | [82 empty] | 1 | [43 empty] | 1
KHV-03 | [82 empty] | 1 | [43 empty] | 1
LAS-UI1 | [54 empty] | 1 | [71 empty] | 1
LAS-UI2 | [54 empty] | 1 | [71 empty] | 1
LAS-UI3 | [35 empty] | 1 | [18 empty] | 2 | [71 empty] | 3
Lift | [9 empty] | 1 | [83 empty] | 1 | [14 empty] | 1 | [8 empty] | 1 | [8 empty] | 4
LIS | [6 empty] | 1 | [119 empty] | 1
LM-01 | [7 empty] | 1 | 1 | [45 empty] | 1 | [19 empty] | 1 | [16 empty] | 1 | [34 empty] | 5
LT-005 | [29 empty] | 1 | [96 empty] | 1
LT-006 | [29 empty] | 1 | [96 empty] | 1
LT-01 | [80 empty] | 1 | [45 empty] | 1
LT-011 | [80 empty] | 1 | [45 empty] | 1
LT-012 | [80 empty] | 1 | [45 empty] | 1
LT-02 | [29 empty] | 1 | [96 empty] | 1
LTL-01 | [80 empty] | 1 | [45 empty] | 1
MC-01 | [29 empty] | 1 | [33 empty] | 1 | [47 empty] | 1 | [14 empty] | 3
MD-08 | [8 empty] | 1 |  |  | 1 | [5 empty] | 1 | [32 empty] | 1 | 1 | [8 empty] | 1 | [22 empty] | 1 | [42 empty] | 7
MD-09 | [32 empty] | 1 | [5 empty] | 1 | [28 empty] | 1 | [20 empty] | 1 | [8 empty] | 1 | [5 empty] | 1 | [7 empty] | 1 | [14 empty] | 7
MDC-01 | 1 | [110 empty] | 1 | [14 empty] | 2
MDC-02 | 1 | [110 empty] | 1 | [14 empty] | 2
MDC-03 | 1 | [110 empty] | 1 | [14 empty] | 2
ML- 01 | [80 empty] | 1 | [45 empty] | 1
ML-01 | [29 empty] | 1 | [96 empty] | 1
ML-02 | [80 empty] | 1 | [45 empty] | 1
MOM | [64 empty] | 1 | [17 empty] | 1 | [43 empty] | 2
MP01 | [80 empty] | 1 | [45 empty] | 1
MPW-01 | 1 | [110 empty] | 1 | [14 empty] | 2
MS-01 | [29 empty] | 1 | [96 empty] | 1
MU-01 | [27 empty] | 1 | [98 empty] | 1
MU-02 | [27 empty] | 1 | [98 empty] | 1
MVT-01 | [8 empty] | 1 | [20 empty] | 1 | [96 empty] | 2
MVT-02 | [29 empty] | 1 | [96 empty] | 1
NAK-02 | [89 empty] | 1 | [36 empty] | 1
NH-01 | [80 empty] | 1 | [45 empty] | 1
NH-02 | [80 empty] | 1 | [45 empty] | 1
NH-03 | [80 empty] | 1 | [45 empty] | 1
NH-04 | [27 empty] | 1 | [98 empty] | 1
NK - 02 | [89 empty] | 1 | [36 empty] | 1
NK - 03 | [89 empty] | 1 | [36 empty] | 1
NK-02 | [28 empty] | 1 | [97 empty] | 1
NK-10 | [89 empty] | 1 | [36 empty] | 1
NK-11 | [28 empty] | 1 | [97 empty] | 1
NK-12 | [89 empty] | 1 | [36 empty] | 1
NK-14 | [89 empty] | 1 | [36 empty] | 1
NKĐT-1 | [28 empty] | 1 | [97 empty] | 1
NKĐT-3 | [28 empty] | 1 | [97 empty] | 1
NKĐT-4 | [28 empty] | 1 | [97 empty] | 1
NKĐT-5 | [28 empty] | 1 | [97 empty] | 1
NKĐT-6 | [28 empty] | 1 | [97 empty] | 1
NKĐT-7 | [28 empty] | 1 | [97 empty] | 1
NKĐT-8 | [28 empty] | 1 | [97 empty] | 1
NU-01 | [27 empty] | 1 | [98 empty] | 1
OVEN-01 |  |  |  | 1 | [60 empty] | 1 | [15 empty] | 1 | [38 empty] | 1 | [6 empty] | 4
PP - 02 | [89 empty] | 1 | [36 empty] | 1
PP - 03 | [89 empty] | 1 | [36 empty] | 1
PP- 04 | [89 empty] | 1 | [36 empty] | 1
PP-002 | [89 empty] | 1 | [36 empty] | 1
PP-006 | [89 empty] | 1 | [36 empty] | 1
PP-008 | [89 empty] | 1 | [36 empty] | 1
PP-02 | [89 empty] | 1 | [36 empty] | 1
PP-03 | [89 empty] | 1 | [36 empty] | 1
PP-04 | [89 empty] | 1 | [36 empty] | 1
PP-05 | [59 empty] | 1 | [29 empty] | 1 | [36 empty] | 2
PP-06 | [28 empty] | 1 | [97 empty] | 1
PP-09 | [32 empty] | 1 | [93 empty] | 1
PP-10 | [28 empty] | 1 | [97 empty] | 1
PP-12 | [89 empty] | 1 | [36 empty] | 1
PP-13 | [28 empty] | 1 | [97 empty] | 1
PP-14 | [28 empty] | 1 | [97 empty] | 1
PP-17 | [28 empty] | 1 | [97 empty] | 1
PP-23 | [28 empty] | 1 | [97 empty] | 1
PP-24 | [28 empty] | 1 | [97 empty] | 1
PT-01 | [59 empty] | 1 | [66 empty] | 1
QS7P-01 | [63 empty] | 1 | 1 | [28 empty] | 1 | [32 empty] | 3
QS7P-02 | [63 empty] | 1 | [31 empty] | 1 | [30 empty] | 2
QSmart | [63 empty] | 1 | [62 empty] | 1
RBM-C8000 | [54 empty] | 1 | [71 empty] | 1
REM-01 | [54 empty] | 1 | [71 empty] | 1
Rida | [78 empty] | 2 | [47 empty] | 2
RM-01 | [59 empty] | 2 |  |  |  |  | 1 | 1 | [21 empty] | 1 | [9 empty] | 1 | [28 empty] | 6
RO-HP | [5 empty] | 1 | [75 empty] | 1 | [18 empty] | 1 | [25 empty] | 3
RO-MB | [31 empty] | 1 | [77 empty] | 1 | [16 empty] | 2
RO-RB | [31 empty] | 1 | [94 empty] | 1
RSM-2 | [37 empty] | 1 | [88 empty] | 1
Spinnaker | 1 | [110 empty] | 1 | [14 empty] | 2
STORE-01 | [19 empty] | 1 |  |  | 1 | [31 empty] | 1 |  |  | 1 | 1 | [23 empty] | 1 |  | 1 | [41 empty] | 7
STR-01 | [33 empty] | 1 | [14 empty] | 1 | [44 empty] | 1 | [28 empty] | 1 |  |  |  | 4
TAC1 | [44 empty] | 1 | [81 empty] | 1
TAC2 | [121 empty] | 1 |  |  |  |  | 1
TAM | [45 empty] | 1 | [12 empty] | 1 | [65 empty] | 1 |  | 3
Tempus600 |  |  | 1 |  | 1 |  | 1 | [7 empty] | 2 | 1 |  |  | 1 |  | 1 | [5 empty] | 1 |  |  | 1 |  |  |  | 1 | [5 empty] | 1 | [12 empty] | 1 | [17 empty] | 1 |  | 1 |  |  |  | 1 | [9 empty] | 1 | [5 empty] | 1 |  |  | 1 | [21 empty] | 1 | 1 |  |  |  |  | 1 |  |  | 22
TH-01 | [28 empty] | 1 | [97 empty] | 1
TH-02 | [28 empty] | 1 | [97 empty] | 1
TH-03 | [28 empty] | 1 | [97 empty] | 1
TH-04 | [28 empty] | 1 | [97 empty] | 1
TI-02 | [89 empty] | 1 | [36 empty] | 1
TL-002 | [80 empty] | 1 | [45 empty] | 1
TL-005 | [27 empty] | 1 | [98 empty] | 1
TL-007 | [80 empty] | 1 | [45 empty] | 1
TL-008 | [80 empty] | 1 | [45 empty] | 1
TL-009 | [27 empty] | 1 | [98 empty] | 1
TL-01 | [80 empty] | 1 | [45 empty] | 1
TL-010 | [80 empty] | 1 | [45 empty] | 1
TL-011 | [80 empty] | 1 | [45 empty] | 1
TL-012 | [80 empty] | 1 | [45 empty] | 1
TL-02 | [80 empty] | 1 | [45 empty] | 1
TL-06 | [27 empty] | 1 | [98 empty] | 1
TL-07 | [27 empty] | 1 | [98 empty] | 1
TOS-01 | [23 empty] | 1 | [11 empty] | 1 | [22 empty] | 1 | [67 empty] | 3
TP-01 | [28 empty] | 1 | [97 empty] | 1
TP-02 | [28 empty] | 1 | [97 empty] | 1
TP-03 | [28 empty] | 1 | [97 empty] | 1
TP-04 | [28 empty] | 1 | [97 empty] | 1
TP-05 | [28 empty] | 1 | [97 empty] | 1
TU-01 | [80 empty] | 1 | [45 empty] | 1
TU-02 | [80 empty] | 1 | [45 empty] | 1
UF-01 | [46 empty] | 1 | [13 empty] | 1 | [18 empty] | 1 | [46 empty] | 3
VT-01 | [29 empty] | 1 | [96 empty] | 1
VT-02 | [29 empty] | 1 | [96 empty] | 1
Grand Total | 13 | 2 | 1 | 1 | 2 | 1 | 2 | 2 | 3 | 1 | 1 | 2 | 1 | 4 | 3 | 2 | 2 | 1 | 2 | 1 | 1 | 1 | 1 | 1 | 2 | 1 | 1 | 15 | 34 | 12 | 2 | 2 | 4 | 2 | 6 | 3 | 4 | 1 | 2 | 2 | 1 | 2 | 1 | 1 | 4 | 5 | 1 | 1 | 5 | 3 | 3 | 2 | 1 | 1 | 12 | 2 | 3 | 2 | 4 | 6 | 2 | 8 | 2 | 5 | 5 | 1 | 2 | 5 | 1 | 1 | 2 | 1 | 5 | 2 | 2 | 1 | 4 | 1 | 4 | 3 | 28 | 2 | 6 | 1 | 3 | 3 | 2 | 2 | 1 | 27 | 2 | 1 | 3 | 6 | 2 | 2 | 1 | 4 | 1 | 2 | 1 | 2 | 1 | 1 | 3 | 4 | 1 | 1 | 2 | 3 | 1 | 16 | 1 | 1 | 2 | 1 | 1 | 3 | 2 | 3 | 1 | 7 | 4 | 1 | 2 | 1 | 426